## Main
Floating potential method | [5 empty] | Ti = 0 | Type into YELLOW cells only |  |  |  |  | R,L in cm, Te in eV, V in volts, I in mA, n in 1011
Rp | L | S | AR | Gas | e | Cs | lD | xp | p | a0 | Vs | Vmin | Vmax |  |  | ABR coeffs: |  | A | B | CC | D | EE | F | G
0.0075 | 1 | 0.047124 | 1.0037 | 40 | 1.6e-19 | 267395.1881 | 0.0021728 | 3.4518 | 1.3333 | 0.60653 | 14.1 | 1.4563 | 13.3562 | Electrons |  |  |  | 0.58318 | 3.7321 | -0.027277 | 5.4306 | 4000.0002 | 7.0112 | 0.095847
Floating potential method Rp: 1. Fill in yellow squares & V-I data in ascending order.; Type into YELLOW cells only p: Iip-V |  |  | -73 | -9.7 | Ions |  | CL coeffs: |  | C1 | C2 | C3 | J | C5 | K
Floating potential method Rp: 2. Vs and the fitting ranges are calculated automatically. |  |  | fitting ranges | [5 empty] | 1278.5937 | 1.2228e-15 | 382.9089 | 244792.4648 | 817772569240.865 | 13.647
Floating potential method Rp: 3. Output data summarized below. | [13 empty] | Estimates
Floating potential method L: Output; Ti = 0 Cs: ni(ABR); Type into YELLOW cells only lD: ni(CL); Type into YELLOW cells only xp: Ion fit | Electron fit |  |  |  |  | Vf and If measurement | [7 empty] | of | Vs | Te
| ABR | CL | Ne | Te | Vs | 2.3234 | 3.4628 | Slope | Intrcpt | Slope | Intrcpt | [5 empty] | hf | [5 empty] | A903 | 14.1 | 2.975
| 2.3234 | 3.4628 | 1.94 | 2.9854 | 14.1 |  |  | -0.020799 | 0.75855 | 0.33496 | -0.97743 |  |  |  |  | Theor= | 4.2552 | Vf |  |  |  |  | 903
Floating potential method Rp: Volts; Floating potential method S: Spline | [5 empty] | Ii at Vf | 0.77575 | 3.9861 | 2.2 |  |  |  |  | -11.3921
| I-(A) | I+(mA) | dI/dV | h | Ii | Iip | Fit Iip | St.line | Ii (fit) | Ie(raw) | Ie | ln Ie | Ie(fit) | I-(A) | Volts | Ii (fit) | h |  | ABR iteration |  |  |  | dI/dV
Floating potential method Rp: -75; Floating potential method L: -0.0026194; Floating potential method Gas: 29.8453; Type into YELLOW cells only xp: 2.3185; Type into YELLOW cells only p: 1.8789 | 1e-05 | 1e-05 |  | 4.6241e-12 | -0.0026194 | -75 | 1.8789 | 29.8453 |  | 3.4628 | 0.0021728 | 3.4518 | 2.5476 | Smooth
Floating potential method Rp: -74.9; Floating potential method L: -0.0022848; Floating potential method Gas: 29.8118; Type into YELLOW cells only xp: 2.3164; Type into YELLOW cells only p: 1.8776 | 1e-05 | 1e-05 |  | 4.7816e-12 | -0.0022848 | -74.9 | 1.8776 | 29.8118 |  | 3.0052 | 0.0023324 | 3.2156 | 2.4666
Floating potential method Rp: -74.8; Floating potential method L: -0.0021361; Floating potential method Gas: 29.7783; Type into YELLOW cells only xp: 2.3143; Type into YELLOW cells only p: 1.8764 | 1e-05 | 1e-05 |  | 4.9445e-12 | -0.0021361 | -74.8 | 1.8764 | 29.7783 |  | 2.7359 | 0.0024445 | 3.0682 | 2.4137
Floating potential method Rp: -74.7; Floating potential method L: -0.0019874; Floating potential method Gas: 29.7448; Type into YELLOW cells only xp: 2.3122; Type into YELLOW cells only p: 1.8751 | 1e-05 | 1e-05 |  | 5.1129e-12 | -0.0019874 | -74.7 | 1.8751 | 29.7448 |  | 2.5748 | 0.0025198 | 2.9765 | 2.3799
Floating potential method Rp: -74.6; Floating potential method L: -0.0019502; Floating potential method Gas: 29.7113; Type into YELLOW cells only xp: 2.3101; Type into YELLOW cells only p: 1.8738 | 1e-05 | 1e-05 |  | 5.2871e-12 | -0.0019502 | -74.6 | 1.8738 | 29.7113 |  | 2.4773 | 0.0025689 | 2.9196 | 2.3585
-74.5 | -0.0019131 | 2.0237 | -0.67347 | 29.6778 | 2.0237 | 2.5598 |  | 2.3081 | 1.8726 | 1e-05 | 1e-05 |  | 5.4672e-12 | -0.0019131 | -74.5 | 1.8726 | 29.6778 |  | 2.4179 | 0.0026002 | 2.8844 | 2.3452
-74.4 | -0.0019131 | 1.9564 | -0.40403 | 29.6443 | 1.9564 | 2.4469 |  | 2.306 | 1.8713 | 1e-05 | 1e-05 |  | 5.6534e-12 | -0.0019131 | -74.4 | 1.8713 | 29.6443 |  | 2.3816 | 0.00262 | 2.8626 | 2.3368
-74.3 | -0.0019131 | 1.916 | -0.26935 | 29.6108 | 1.916 | 2.3797 |  | 2.3039 | 1.87 | 1e-05 | 1e-05 |  | 5.846e-12 | -0.0019131 | -74.3 | 1.87 | 29.6108 |  | 2.3592 | 0.0026324 | 2.8491 | 2.3317
-74.2 | -0.0018758 | 1.8891 | -0.10108 | 29.5773 | 1.8891 | 2.3352 |  | 2.3018 | 1.8688 | 1e-05 | 1e-05 |  | 6.0451e-12 | -0.0018758 | -74.2 | 1.8688 | 29.5773 |  | 2.3454 | 0.0026401 | 2.8408 | 2.3285
-74.1 | -0.0018758 | 1.8789 | -0.067384 | 29.5438 | 1.8789 | 2.3186 |  | 2.2997 | 1.8675 | 1e-05 | 1e-05 |  | 6.251e-12 | -0.0018758 | -74.1 | 1.8675 | 29.5438 |  | 2.337 | 0.0026449 | 2.8357 | 2.3265
-74 | -0.0018758 | 1.8722 | -0.067384 | 29.5103 | 1.8722 | 2.3075 |  | 2.2977 | 1.8662 | 1e-05 | 1e-05 |  | 6.464e-12 | -0.0018758 | -74 | 1.8662 | 29.5103 |  | 2.3317 | 0.0026479 | 2.8325 | 2.3253 | -0.15613
-73.9 | -0.0018758 | 1.8655 | -0.033782 | 29.4768 | 1.8655 | 2.2964 |  | 2.2956 | 1.865 | 1e-05 | 1e-05 |  | 6.6841e-12 | -0.0018758 | -73.9 | 1.865 | 29.4768 |  | 2.3285 | 0.0026497 | 2.8305 | 2.3245 | -0.097963
-73.8 | -0.0018387 | 1.8621 | -0.033782 | 29.4433 | 1.8621 | 2.2909 |  | 2.2935 | 1.8637 | 1e-05 | 0.031865 |  | 6.9118e-12 | -0.0018387 | -73.8 | 1.8637 | 29.4433 |  | 2.3265 | 0.0026508 | 2.8293 | 2.324 | -0.061233
-73.7 | -0.0018387 | 1.8587 | -0.033601 | 29.4099 | 1.8587 | 2.2853 |  | 2.2914 | 1.8624 | 1e-05 | 0.030597 |  | 7.1473e-12 | -0.0018387 | -73.7 | 1.8624 | 29.4099 |  | 2.3253 | 0.0026515 | 2.8286 | 2.3237 | -0.036746
-73.6 | -0.0018758 | 1.8554 | -0.033601 | 29.3764 | 1.8554 | 2.2798 |  | 2.2893 | 1.8612 | 1e-05 | 1e-05 |  | 7.3907e-12 | -0.0018758 | -73.6 | 1.8612 | 29.3764 |  | 2.3245 | 0.002652 | 2.8281 | 2.3236 | -0.030612
-73.5 | -0.0018758 | 1.852 | 0 | 29.3429 | 1.852 | 2.2743 |  | 2.2873 | 1.8599 | 1e-05 | 1e-05 |  | 7.6425e-12 | -0.0018758 | -73.5 | 1.8599 | 29.3429 |  | 2.324 | 0.0026522 | 2.8278 | 2.3235 | -0.027541
-73.4 | -0.0018387 | 1.852 | -0.033601 | 29.3094 | 1.852 | 2.2743 |  | 2.2852 | 1.8586 | 1e-05 | 0.026792 |  | 7.9028e-12 | -0.0018387 | -73.4 | 1.8586 | 29.3094 |  | 2.3238 | 0.0026524 | 2.8276 | 2.3234 | -0.018361
-73.3 | -0.0018758 | 1.8486 | 0 | 29.2759 | 1.8486 | 2.2688 |  | 2.2831 | 1.8574 | 1e-05 | 1e-05 |  | 8.172e-12 | -0.0018758 | -73.3 | 1.8574 | 29.2759 |  | 2.3236 | 0.0026525 | 2.8275 | 2.3234 | -0.01529
-73.2 | -0.0018758 | 1.8486 | 0 | 29.2424 | 1.8486 | 2.2688 |  | 2.281 | 1.8561 | 1e-05 | 1e-05 |  | 8.4504e-12 | -0.0018758 | -73.2 | 1.8561 | 29.2424 |  | 2.3235 | 0.0026526 | 2.8275 | 2.3233 | -0.015273
-73.1 | -0.0018387 | 1.8486 | -0.033601 | 29.2089 | 1.8486 | 2.2688 |  | 2.2789 | 1.8548 | 1e-05 | 0.022985 |  | 8.7382e-12 | -0.0018387 | -73.1 | 1.8548 | 29.2089 |  | 2.3234 | 0.0026526 | 2.8274 | 2.3233 | -0.012219
-73 | -0.0018387 | 1.8453 | -0.033601 | 29.1754 | 1.8453 | 2.2633 | 2.2633 | 2.2769 | 1.8535 | 1e-05 | 0.021715 |  | 9.0359e-12 | -0.0018387 | -73 | 1.8535 | 29.1754 | [5 empty] | -0.009164
-72.9 | -0.0018758 | 1.8419 | 0.033601 | 29.1419 | 1.8419 | 2.2578 | 2.2578 | 2.2748 | 1.8523 | 1e-05 | 1e-05 |  | 9.3437e-12 | -0.0018758 | -72.9 | 1.8523 | 29.1419 | [5 empty] | -0.012219
-72.8 | -0.0018387 | 1.8453 | 0 | 29.1084 | 1.8453 | 2.2633 | 2.2633 | 2.2727 | 1.851 | 1e-05 | 0.019175 |  | 9.662e-12 | -0.0018387 | -72.8 | 1.851 | 29.1084 | [5 empty] | -0.009164
-72.7 | -0.0018387 | 1.8453 | -0.033601 | 29.0749 | 1.8453 | 2.2633 | 2.2633 | 2.2706 | 1.8497 | 1e-05 | 0.017904 |  | 9.9911e-12 | -0.0018387 | -72.7 | 1.8497 | 29.0749 | [5 empty] | -0.009164
-72.6 | -0.0018387 | 1.8419 | 0 | 29.0414 | 1.8419 | 2.2578 | 2.2578 | 2.2685 | 1.8485 | 1e-05 | 0.016633 |  | 1.0331e-11 | -0.0018387 | -72.6 | 1.8485 | 29.0414 | [5 empty] | -0.012219
-72.5 | -0.0018387 | 1.8419 | 0 | 29.0079 | 1.8419 | 2.2578 | 2.2578 | 2.2665 | 1.8472 | 1e-05 | 0.015362 |  | 1.0683e-11 | -0.0018387 | -72.5 | 1.8472 | 29.0079 | [5 empty] | -0.009164
-72.4 | -0.0018387 | 1.8419 | -0.033601 | 28.9744 | 1.8419 | 2.2578 | 2.2578 | 2.2644 | 1.8459 | 1e-05 | 0.01409 |  | 1.1047e-11 | -0.0018387 | -72.4 | 1.8459 | 28.9744 | [5 empty] | -0.009164
-72.3 | -0.0018758 | 1.8386 | 0 | 28.9409 | 1.8386 | 2.2523 | 2.2523 | 2.2623 | 1.8446 | 1e-05 | 1e-05 |  | 1.1424e-11 | -0.0018758 | -72.3 | 1.8446 | 28.9409 | [5 empty] | -0.015273
-72.2 | -0.0018758 | 1.8386 | 0 | 28.9074 | 1.8386 | 2.2523 | 2.2523 | 2.2602 | 1.8434 | 1e-05 | 1e-05 |  | 1.1813e-11 | -0.0018758 | -72.2 | 1.8434 | 28.9074 | [5 empty] | -0.018328
-72.1 | -0.0018387 | 1.8386 | -0.033601 | 28.8739 | 1.8386 | 2.2523 | 2.2523 | 2.2581 | 1.8421 | 1e-05 | 0.010274 |  | 1.2215e-11 | -0.0018387 | -72.1 | 1.8421 | 28.8739 | [5 empty] | -0.015273
-72 | -0.0018387 | 1.8352 | 0 | 28.8404 | 1.8352 | 2.2469 | 2.2469 | 2.2561 | 1.8408 | 1e-05 | 0.0090016 |  | 1.2631e-11 | -0.0018387 | -72 | 1.8408 | 28.8404 | [5 empty] | -0.015273
-71.9 | -0.0018387 | 1.8352 | -0.033601 | 28.8069 | 1.8352 | 2.2469 | 2.2469 | 2.254 | 1.8396 | 1e-05 | 0.0077286 |  | 1.3061e-11 | -0.0018387 | -71.9 | 1.8396 | 28.8069 | [5 empty] | -0.018328
-71.8 | -0.0018387 | 1.8318 | -0.033601 | 28.7734 | 1.8318 | 2.2414 | 2.2414 | 2.2519 | 1.8383 | 1e-05 | 0.0064553 |  | 1.3506e-11 | -0.0018387 | -71.8 | 1.8383 | 28.7734 | [5 empty] | -0.018328
-71.7 | -0.0018387 | 1.8285 | -0.033601 | 28.7399 | 1.8285 | 2.2359 | 2.2359 | 2.2498 | 1.837 | 1e-05 | 0.0051818 |  | 1.3966e-11 | -0.0018387 | -71.7 | 1.837 | 28.7399 | [5 empty] | -0.018328
-71.6 | -0.0018387 | 1.8251 | 0 | 28.7064 | 1.8251 | 2.2304 | 2.2304 | 2.2477 | 1.8357 | 1e-05 | 0.003908 |  | 1.4442e-11 | -0.0018387 | -71.6 | 1.8357 | 28.7064 | [5 empty] | -0.021383
-71.5 | -0.0018016 | 1.8251 | 0 | 28.6729 | 1.8251 | 2.2304 | 2.2304 | 2.2457 | 1.8345 | 1e-05 | 0.039595 |  | 1.4934e-11 | -0.0018016 | -71.5 | 1.8345 | 28.6729 | [5 empty] | -0.015273
-71.4 | -0.0018387 | 1.8251 | -0.033601 | 28.6394 | 1.8251 | 2.2304 | 2.2304 | 2.2436 | 1.8332 | 1e-05 | 0.0013594 |  | 1.5443e-11 | -0.0018387 | -71.4 | 1.8332 | 28.6394 | [5 empty] | -0.018328
-71.3 | -0.0018016 | 1.8218 | -0.033601 | 28.6059 | 1.8218 | 2.2249 | 2.2249 | 2.2415 | 1.8319 | 1e-05 | 0.037046 |  | 1.5969e-11 | -0.0018016 | -71.3 | 1.8319 | 28.6059 | [5 empty] | -0.015273
-71.2 | -0.0018387 | 1.8184 | 0 | 28.5724 | 1.8184 | 2.2195 | 2.2195 | 2.2394 | 1.8306 | 1e-05 | 1e-05 |  | 1.6513e-11 | -0.0018387 | -71.2 | 1.8306 | 28.5724 | [5 empty] | -0.015273
-71.1 | -0.0018387 | 1.8184 | -0.033601 | 28.5389 | 1.8184 | 2.2195 | 2.2195 | 2.2373 | 1.8294 | 1e-05 | 1e-05 |  | 1.7075e-11 | -0.0018387 | -71.1 | 1.8294 | 28.5389 | [5 empty] | -0.012219
-71 | -0.0018387 | 1.815 | 0.033601 | 28.5055 | 1.815 | 2.214 | 2.214 | 2.2353 | 1.8281 | 1e-05 | 1e-05 |  | 1.7657e-11 | -0.0018387 | -71 | 1.8281 | 28.5055 | [5 empty] | -0.015273
-70.9 | -0.0018387 | 1.8184 | -0.033601 | 28.472 | 1.8184 | 2.2195 | 2.2195 | 2.2332 | 1.8268 | 1e-05 | 1e-05 |  | 1.8258e-11 | -0.0018387 | -70.9 | 1.8268 | 28.472 | [5 empty] | -0.018328
-70.8 | -0.0018016 | 1.815 | 0 | 28.4385 | 1.815 | 2.214 | 2.214 | 2.2311 | 1.8255 | 1e-05 | 0.030668 |  | 1.888e-11 | -0.0018016 | -70.8 | 1.8255 | 28.4385 | [5 empty] | -0.015273
-70.7 | -0.0018016 | 1.815 | -0.033601 | 28.405 | 1.815 | 2.214 | 2.214 | 2.229 | 1.8243 | 1e-05 | 0.029391 |  | 1.9524e-11 | -0.0018016 | -70.7 | 1.8243 | 28.405 | [5 empty] | -0.012219
-70.6 | -0.0018387 | 1.8117 | 0 | 28.3715 | 1.8117 | 2.2085 | 2.2085 | 2.2269 | 1.823 | 1e-05 | 1e-05 |  | 2.0189e-11 | -0.0018387 | -70.6 | 1.823 | 28.3715 | [5 empty] | -0.015273
-70.5 | -0.0018016 | 1.8117 | -0.033601 | 28.338 | 1.8117 | 2.2085 | 2.2085 | 2.2249 | 1.8217 | 1e-05 | 0.026838 |  | 2.0876e-11 | -0.0018016 | -70.5 | 1.8217 | 28.338 | [5 empty] | -0.012219
-70.4 | -0.0018387 | 1.8083 | -0.033601 | 28.3045 | 1.8083 | 2.2031 | 2.2031 | 2.2228 | 1.8204 | 1e-05 | 1e-05 |  | 2.1587e-11 | -0.0018387 | -70.4 | 1.8204 | 28.3045 | [5 empty] | -0.018328
-70.3 | -0.0018016 | 1.8049 | 0 | 28.271 | 1.8049 | 2.1976 | 2.1976 | 2.2207 | 1.8192 | 1e-05 | 0.024282 |  | 2.2323e-11 | -0.0018016 | -70.3 | 1.8192 | 28.271 | [5 empty] | -0.015273
-70.2 | -0.0018016 | 1.8049 | 0 | 28.2375 | 1.8049 | 2.1976 | 2.1976 | 2.2186 | 1.8179 | 1e-05 | 0.023004 |  | 2.3083e-11 | -0.0018016 | -70.2 | 1.8179 | 28.2375 | [5 empty] | -0.015273
-70.1 | -0.0018016 | 1.8049 | -0.033601 | 28.204 | 1.8049 | 2.1976 | 2.1976 | 2.2166 | 1.8166 | 1e-05 | 0.021726 |  | 2.3869e-11 | -0.0018016 | -70.1 | 1.8166 | 28.204 | [5 empty] | -0.012219
-70 | -0.0018387 | 1.8016 | 0 | 28.1705 | 1.8016 | 2.1922 | 2.1922 | 2.2145 | 1.8153 | 1e-05 | 1e-05 |  | 2.4683e-11 | -0.0018387 | -70 | 1.8153 | 28.1705 | [5 empty] | -0.015273
-69.9 | -0.0018016 | 1.8016 | -0.033601 | 28.137 | 1.8016 | 2.1922 | 2.1922 | 2.2124 | 1.814 | 1e-05 | 0.019169 |  | 2.5523e-11 | -0.0018016 | -69.9 | 1.814 | 28.137 | [5 empty] | -0.012219
-69.8 | -0.0018016 | 1.7982 | 0 | 28.1035 | 1.7982 | 2.1867 | 2.1867 | 2.2103 | 1.8128 | 1e-05 | 0.017889 |  | 2.6393e-11 | -0.0018016 | -69.8 | 1.8128 | 28.1035 | [5 empty] | -0.012235
-69.7 | -0.0018016 | 1.7982 | 0 | 28.07 | 1.7982 | 2.1867 | 2.1867 | 2.2082 | 1.8115 | 1e-05 | 0.01661 |  | 2.7292e-11 | -0.0018016 | -69.7 | 1.8115 | 28.07 | [5 empty] | -0.015306
-69.6 | -0.0018016 | 1.7982 | 0 | 28.0365 | 1.7982 | 2.1867 | 2.1867 | 2.2062 | 1.8102 | 1e-05 | 0.01533 |  | 2.8221e-11 | -0.0018016 | -69.6 | 1.8102 | 28.0365 | [5 empty] | -0.015306
-69.5 | -0.0018016 | 1.7982 | -0.033601 | 28.003 | 1.7982 | 2.1867 | 2.1867 | 2.2041 | 1.8089 | 1e-05 | 0.01405 |  | 2.9183e-11 | -0.0018016 | -69.5 | 1.8089 | 28.003 | [5 empty] | -0.012252
-69.4 | -0.0018016 | 1.7949 | 0 | 27.9695 | 1.7949 | 2.1813 | 2.1813 | 2.202 | 1.8076 | 1e-05 | 0.01277 |  | 3.0177e-11 | -0.0018016 | -69.4 | 1.8076 | 27.9695 | [5 empty] | -0.012252
-69.3 | -0.0018016 | 1.7949 | -0.033782 | 27.936 | 1.7949 | 2.1813 | 2.1813 | 2.1999 | 1.8064 | 1e-05 | 0.011489 |  | 3.1205e-11 | -0.0018016 | -69.3 | 1.8064 | 27.936 | [5 empty] | -0.012268
-69.2 | -0.0018016 | 1.7915 | -0.033782 | 27.9025 | 1.7915 | 2.1758 | 2.1758 | 2.1978 | 1.8051 | 1e-05 | 0.010208 |  | 3.2268e-11 | -0.0018016 | -69.2 | 1.8051 | 27.9025 | [5 empty] | -0.012268
-69.1 | -0.0018016 | 1.7881 | 0 | 27.869 | 1.7881 | 2.1703 | 2.1703 | 2.1958 | 1.8038 | 1e-05 | 0.0089267 |  | 3.3367e-11 | -0.0018016 | -69.1 | 1.8038 | 27.869 | [5 empty] | -0.015339
-69 | -0.0018016 | 1.7881 | 0 | 27.8355 | 1.7881 | 2.1703 | 2.1703 | 2.1937 | 1.8025 | 1e-05 | 0.007645 |  | 3.4504e-11 | -0.0018016 | -69 | 1.8025 | 27.8355 | [5 empty] | -0.01841
-68.9 | -0.0018016 | 1.7881 | 0 | 27.802 | 1.7881 | 2.1703 | 2.1703 | 2.1916 | 1.8012 | 1e-05 | 0.0063631 |  | 3.5679e-11 | -0.0018016 | -68.9 | 1.8012 | 27.802 | [5 empty] | -0.018427
-68.8 | -0.0018016 | 1.7881 | -0.033782 | 27.7685 | 1.7881 | 2.1703 | 2.1703 | 2.1895 | 1.8 | 1e-05 | 0.0050809 |  | 3.6894e-11 | -0.0018016 | -68.8 | 1.8 | 27.7685 | [5 empty] | -0.021498
-68.7 | -0.0017643 | 1.7847 | 0 | 27.735 | 1.7847 | 2.1649 | 2.1649 | 2.1874 | 1.7987 | 1e-05 | 0.040959 |  | 3.8151e-11 | -0.0017643 | -68.7 | 1.7987 | 27.735 | [5 empty] | -0.018427
-68.6 | -0.0017643 | 1.7847 | -0.033782 | 27.7015 | 1.7847 | 2.1649 | 2.1649 | 2.1854 | 1.7974 | 1e-05 | 0.039676 |  | 3.9451e-11 | -0.0017643 | -68.6 | 1.7974 | 27.7015 | [5 empty] | -0.015356
-68.5 | -0.0018016 | 1.7814 | -0.033782 | 27.668 | 1.7814 | 2.1594 | 2.1594 | 2.1833 | 1.7961 | 1e-05 | 0.0012324 |  | 4.0794e-11 | -0.0018016 | -68.5 | 1.7961 | 27.668 | [5 empty] | -0.018427
-68.4 | -0.0018016 | 1.778 | -0.033782 | 27.6345 | 1.778 | 2.1539 | 2.1539 | 2.1812 | 1.7948 | 1e-05 | 1e-05 |  | 4.2184e-11 | -0.0018016 | -68.4 | 1.7948 | 27.6345 | [5 empty] | -0.021498
-68.3 | -0.0018016 | 1.7746 | -0.033782 | 27.601 | 1.7746 | 2.1485 | 2.1485 | 2.1791 | 1.7935 | 1e-05 | 1e-05 |  | 4.3621e-11 | -0.0018016 | -68.3 | 1.7935 | 27.601 | [5 empty] | -0.021498
-68.2 | -0.0017643 | 1.7712 | 0 | 27.5676 | 1.7712 | 2.143 | 2.143 | 2.177 | 1.7923 | 1e-05 | 0.034542 |  | 4.5107e-11 | -0.0017643 | -68.2 | 1.7923 | 27.5676 | [5 empty] | -0.015356
-68.1 | -0.0018016 | 1.7712 | 0 | 27.5341 | 1.7712 | 2.143 | 2.143 | 2.175 | 1.791 | 1e-05 | 1e-05 |  | 4.6643e-11 | -0.0018016 | -68.1 | 1.791 | 27.5341 | [5 empty] | -0.018427
-68 | -0.0017643 | 1.7712 | -0.033782 | 27.5006 | 1.7712 | 2.143 | 2.143 | 2.1729 | 1.7897 | 1e-05 | 0.031973 |  | 4.8232e-11 | -0.0017643 | -68 | 1.7897 | 27.5006 | [5 empty] | -0.015356
-67.9 | -0.0017643 | 1.7678 | -0.033782 | 27.4671 | 1.7678 | 2.1376 | 2.1376 | 2.1708 | 1.7884 | 1e-05 | 0.030688 |  | 4.9875e-11 | -0.0017643 | -67.9 | 1.7884 | 27.4671 | [5 empty] | -0.012285
-67.8 | -0.0017643 | 1.7645 | 0 | 27.4336 | 1.7645 | 2.1321 | 2.1321 | 2.1687 | 1.7871 | 1e-05 | 0.029402 |  | 5.1574e-11 | -0.0017643 | -67.8 | 1.7871 | 27.4336 | [5 empty] | -0.0092134
-67.7 | -0.0017643 | 1.7645 | 0.033782 | 27.4001 | 1.7645 | 2.1321 | 2.1321 | 2.1666 | 1.7858 | 1e-05 | 0.028117 |  | 5.3331e-11 | -0.0017643 | -67.7 | 1.7858 | 27.4001 | [5 empty] | -0.0030711
-67.6 | -0.0017643 | 1.7678 | -0.033782 | 27.3666 | 1.7678 | 2.1376 | 2.1376 | 2.1646 | 1.7845 | 1e-05 | 0.026831 |  | 5.5148e-11 | -0.0017643 | -67.6 | 1.7845 | 27.3666 | [5 empty] | -0.0030711
-67.5 | -0.0017643 | 1.7645 | 0 | 27.3331 | 1.7645 | 2.1321 | 2.1321 | 2.1625 | 1.7833 | 1e-05 | 0.025545 |  | 5.7026e-11 | -0.0017643 | -67.5 | 1.7833 | 27.3331 | [5 empty] | -0.0030711
-67.4 | -0.0017643 | 1.7645 | 0 | 27.2996 | 1.7645 | 2.1321 | 2.1321 | 2.1604 | 1.782 | 1e-05 | 0.024258 |  | 5.8969e-11 | -0.0017643 | -67.4 | 1.782 | 27.2996 | [5 empty] | -4.3652e-16
-67.3 | -0.0017643 | 1.7645 | 0 | 27.2661 | 1.7645 | 2.1321 | 2.1321 | 2.1583 | 1.7807 | 1e-05 | 0.022971 |  | 6.0977e-11 | -0.0017643 | -67.3 | 1.7807 | 27.2661 | [5 empty] | 0.0030711
-67.2 | -0.0018016 | 1.7645 | 0.033782 | 27.2326 | 1.7645 | 2.1321 | 2.1321 | 2.1562 | 1.7794 | 1e-05 | 1e-05 |  | 6.3054e-11 | -0.0018016 | -67.2 | 1.7794 | 27.2326 | [5 empty] | -4.3652e-16
-67.1 | -0.0018016 | 1.7678 | 0 | 27.1991 | 1.7678 | 2.1376 | 2.1376 | 2.1542 | 1.7781 | 1e-05 | 1e-05 |  | 6.5202e-11 | -0.0018016 | -67.1 | 1.7781 | 27.1991 | [5 empty] | -0.0061423
-67 | -0.0017643 | 1.7678 | 0 | 27.1656 | 1.7678 | 2.1376 | 2.1376 | 2.1521 | 1.7768 | 1e-05 | 0.019109 |  | 6.7423e-11 | -0.0017643 | -67 | 1.7768 | 27.1656 | [5 empty] | -0.0030711
-66.9 | -0.0017643 | 1.7678 | 0 | 27.1321 | 1.7678 | 2.1376 | 2.1376 | 2.15 | 1.7755 | 1e-05 | 0.017821 |  | 6.972e-11 | -0.0017643 | -66.9 | 1.7755 | 27.1321 | [5 empty] | -0.0061258
-66.8 | -0.0017643 | 1.7678 | 0 | 27.0986 | 1.7678 | 2.1376 | 2.1376 | 2.1479 | 1.7742 | 1e-05 | 0.016533 |  | 7.2095e-11 | -0.0017643 | -66.8 | 1.7742 | 27.0986 | [5 empty] | -0.0091804
-66.7 | -0.0017643 | 1.7678 | -0.033782 | 27.0651 | 1.7678 | 2.1376 | 2.1376 | 2.1458 | 1.773 | 1e-05 | 0.015244 |  | 7.4551e-11 | -0.0017643 | -66.7 | 1.773 | 27.0651 | [5 empty] | -0.012235
-66.6 | -0.0018016 | 1.7645 | -0.033782 | 27.0316 | 1.7645 | 2.1321 | 2.1321 | 2.1438 | 1.7717 | 1e-05 | 1e-05 |  | 7.709e-11 | -0.0018016 | -66.6 | 1.7717 | 27.0316 | [5 empty] | -0.021432
-66.5 | -0.0017643 | 1.7611 | 0 | 26.9981 | 1.7611 | 2.1267 | 2.1267 | 2.1417 | 1.7704 | 1e-05 | 0.012666 |  | 7.9716e-11 | -0.0017643 | -66.5 | 1.7704 | 26.9981 | [5 empty] | -0.021432
-66.4 | -0.0017643 | 1.7611 | -0.033601 | 26.9646 | 1.7611 | 2.1267 | 2.1267 | 2.1396 | 1.7691 | 1e-05 | 0.011376 |  | 8.2432e-11 | -0.0017643 | -66.4 | 1.7691 | 26.9646 | [5 empty] | -0.021432
-66.3 | -0.0017643 | 1.7577 | -0.033601 | 26.9311 | 1.7577 | 2.1213 | 2.1213 | 2.1375 | 1.7678 | 1e-05 | 0.010086 |  | 8.524e-11 | -0.0017643 | -66.3 | 1.7678 | 26.9311 | [5 empty] | -0.021432
-66.2 | -0.0017643 | 1.7544 | -0.033601 | 26.8976 | 1.7544 | 2.1159 | 2.1159 | 2.1354 | 1.7665 | 1e-05 | 0.0087957 |  | 8.8143e-11 | -0.0017643 | -66.2 | 1.7665 | 26.8976 | [5 empty] | -0.024487
-66.1 | -0.0017643 | 1.751 | -0.067384 | 26.8641 | 1.751 | 2.1105 | 2.1105 | 2.1334 | 1.7652 | 1e-05 | 0.0075052 |  | 9.1146e-11 | -0.0017643 | -66.1 | 1.7652 | 26.8641 | [5 empty] | -0.021416
-66 | -0.0017643 | 1.7443 | 0 | 26.8306 | 1.7443 | 2.0997 | 2.0997 | 2.1313 | 1.7639 | 1e-05 | 0.0062143 |  | 9.425e-11 | -0.0017643 | -66 | 1.7639 | 26.8306 | [5 empty] | -0.018344
-65.9 | -0.0017643 | 1.7443 | 0 | 26.7971 | 1.7443 | 2.0997 | 2.0997 | 2.1292 | 1.7626 | 1e-05 | 0.0049231 |  | 9.7461e-11 | -0.0017643 | -65.9 | 1.7626 | 26.7971 | [5 empty] | -0.018344
-65.8 | -0.0017272 | 1.7443 | 0 | 26.7636 | 1.7443 | 2.0997 | 2.0997 | 2.1271 | 1.7613 | 1e-05 | 0.040593 |  | 1.0078e-10 | -0.0017272 | -65.8 | 1.7613 | 26.7636 | [5 empty] | -0.01529
-65.7 | -0.0017272 | 1.7443 | -0.033601 | 26.7301 | 1.7443 | 2.0997 | 2.0997 | 2.125 | 1.76 | 1e-05 | 0.039301 |  | 1.0421e-10 | -0.0017272 | -65.7 | 1.76 | 26.7301 | [5 empty] | -0.012235
-65.6 | -0.0017272 | 1.7409 | 0 | 26.6966 | 1.7409 | 2.0943 | 2.0943 | 2.123 | 1.7588 | 1e-05 | 0.038009 |  | 1.0776e-10 | -0.0017272 | -65.6 | 1.7588 | 26.6966 | [5 empty] | -0.0061258
-65.5 | -0.0017272 | 1.7409 | 0 | 26.6631 | 1.7409 | 2.0943 | 2.0943 | 2.1209 | 1.7575 | 1e-05 | 0.036716 |  | 1.1143e-10 | -0.0017272 | -65.5 | 1.7575 | 26.6631 | [5 empty] | 4.34e-16
-65.4 | -0.0017643 | 1.7409 | 0 | 26.6297 | 1.7409 | 2.0943 | 2.0943 | 2.1188 | 1.7562 | 1e-05 | 1e-05 |  | 1.1523e-10 | -0.0017643 | -65.4 | 1.7562 | 26.6297 | [5 empty] | 4.34e-16
-65.3 | -0.0017643 | 1.7409 | 0 | 26.5962 | 1.7409 | 2.0943 | 2.0943 | 2.1167 | 1.7549 | 1e-05 | 1e-05 |  | 1.1916e-10 | -0.0017643 | -65.3 | 1.7549 | 26.5962 | [5 empty] | 4.34e-16
-65.2 | -0.0017643 | 1.7409 | 0 | 26.5627 | 1.7409 | 2.0943 | 2.0943 | 2.1146 | 1.7536 | 1e-05 | 1e-05 |  | 1.2321e-10 | -0.0017643 | -65.2 | 1.7536 | 26.5627 | [5 empty] | -0.0030547
-65.1 | -0.0017272 | 1.7409 | 0.033601 | 26.5292 | 1.7409 | 2.0943 | 2.0943 | 2.1126 | 1.7523 | 1e-05 | 0.031543 |  | 1.2741e-10 | -0.0017272 | -65.1 | 1.7523 | 26.5292 | [5 empty] | 0.0030547
-65 | -0.0017643 | 1.7443 | 0 | 26.4957 | 1.7443 | 2.0997 | 2.0997 | 2.1105 | 1.751 | 1e-05 | 1e-05 |  | 1.3175e-10 | -0.0017643 | -65 | 1.751 | 26.4957 | [5 empty] | 4.34e-16
-64.9 | -0.0017643 | 1.7443 | 0 | 26.4622 | 1.7443 | 2.0997 | 2.0997 | 2.1084 | 1.7497 | 1e-05 | 1e-05 |  | 1.3624e-10 | -0.0017643 | -64.9 | 1.7497 | 26.4622 | [5 empty] | -0.0030547
-64.8 | -0.0017643 | 1.7443 | 0 | 26.4287 | 1.7443 | 2.0997 | 2.0997 | 2.1063 | 1.7484 | 1e-05 | 1e-05 |  | 1.4088e-10 | -0.0017643 | -64.8 | 1.7484 | 26.4287 | [5 empty] | -0.0030547
-64.7 | -0.0017272 | 1.7443 | -0.033601 | 26.3952 | 1.7443 | 2.0997 | 2.0997 | 2.1042 | 1.7471 | 1e-05 | 0.026365 |  | 1.4568e-10 | -0.0017272 | -64.7 | 1.7471 | 26.3952 | [5 empty] | -0.0030547
-64.6 | -0.0017272 | 1.7409 | 0.033601 | 26.3617 | 1.7409 | 2.0943 | 2.0943 | 2.1022 | 1.7458 | 1e-05 | 0.02507 |  | 1.5064e-10 | -0.0017272 | -64.6 | 1.7458 | 26.3617 | [5 empty] | -0.0030547
-64.5 | -0.0017643 | 1.7443 | -0.033601 | 26.3282 | 1.7443 | 2.0997 | 2.0997 | 2.1001 | 1.7445 | 1e-05 | 1e-05 |  | 1.5577e-10 | -0.0017643 | -64.5 | 1.7445 | 26.3282 | [5 empty] | -0.009164
-64.4 | -0.0017272 | 1.7409 | -0.033601 | 26.2947 | 1.7409 | 2.0943 | 2.0943 | 2.098 | 1.7432 | 1e-05 | 0.022478 |  | 1.6108e-10 | -0.0017272 | -64.4 | 1.7432 | 26.2947 | [5 empty] | -0.009164
-64.3 | -0.0017643 | 1.7375 | 0 | 26.2612 | 1.7375 | 2.0889 | 2.0889 | 2.0959 | 1.7419 | 1e-05 | 1e-05 |  | 1.6657e-10 | -0.0017643 | -64.3 | 1.7419 | 26.2612 | [5 empty] | -0.012219
-64.2 | -0.0017643 | 1.7375 | 0 | 26.2277 | 1.7375 | 2.0889 | 2.0889 | 2.0938 | 1.7406 | 1e-05 | 1e-05 |  | 1.7224e-10 | -0.0017643 | -64.2 | 1.7406 | 26.2277 | [5 empty] | -0.015273
-64.1 | -0.0017272 | 1.7375 | 0 | 26.1942 | 1.7375 | 2.0889 | 2.0889 | 2.0918 | 1.7393 | 1e-05 | 0.018588 |  | 1.7811e-10 | -0.0017272 | -64.1 | 1.7393 | 26.1942 | [5 empty] | -0.012219
-64 | -0.0017643 | 1.7375 | -0.033601 | 26.1607 | 1.7375 | 2.0889 | 2.0889 | 2.0897 | 1.738 | 1e-05 | 1e-05 |  | 1.8417e-10 | -0.0017643 | -64 | 1.738 | 26.1607 | [5 empty] | -0.018328
-63.9 | -0.0017272 | 1.7342 | 0 | 26.1272 | 1.7342 | 2.0835 | 2.0835 | 2.0876 | 1.7367 | 1e-05 | 0.015994 |  | 1.9045e-10 | -0.0017272 | -63.9 | 1.7367 | 26.1272 | [5 empty] | -0.015273
-63.8 | -0.0017272 | 1.7342 | -0.033601 | 26.0937 | 1.7342 | 2.0835 | 2.0835 | 2.0855 | 1.7354 | 1e-05 | 0.014696 |  | 1.9694e-10 | -0.0017272 | -63.8 | 1.7354 | 26.0937 | [5 empty] | -0.012219
-63.7 | -0.0017643 | 1.7308 | -0.033601 | 26.0602 | 1.7308 | 2.0781 | 2.0781 | 2.0834 | 1.7341 | 1e-05 | 1e-05 |  | 2.0364e-10 | -0.0017643 | -63.7 | 1.7341 | 26.0602 | [5 empty] | -0.015273
-63.6 | -0.0017272 | 1.7275 | 0 | 26.0267 | 1.7275 | 2.0727 | 2.0727 | 2.0814 | 1.7328 | 1e-05 | 0.012099 |  | 2.1058e-10 | -0.0017272 | -63.6 | 1.7328 | 26.0267 | [5 empty] | -0.015273
-63.5 | -0.0017272 | 1.7275 | -0.033601 | 25.9932 | 1.7275 | 2.0727 | 2.0727 | 2.0793 | 1.7315 | 1e-05 | 0.0108 |  | 2.1775e-10 | -0.0017272 | -63.5 | 1.7315 | 25.9932 | [5 empty] | -0.015273
-63.4 | -0.0017272 | 1.7241 | 0 | 25.9597 | 1.7241 | 2.0674 | 2.0674 | 2.0772 | 1.7302 | 1e-05 | 0.0095008 |  | 2.2517e-10 | -0.0017272 | -63.4 | 1.7302 | 25.9597 | [5 empty] | -0.012219
-63.3 | -0.0017272 | 1.7241 | 0 | 25.9262 | 1.7241 | 2.0674 | 2.0674 | 2.0751 | 1.7289 | 1e-05 | 0.0082013 |  | 2.3284e-10 | -0.0017272 | -63.3 | 1.7289 | 25.9262 | [5 empty] | -0.012219
-63.2 | -0.0017272 | 1.7241 | -0.033601 | 25.8927 | 1.7241 | 2.0674 | 2.0674 | 2.073 | 1.7276 | 1e-05 | 0.0069014 |  | 2.4077e-10 | -0.0017272 | -63.2 | 1.7276 | 25.8927 | [5 empty] | -0.012227
-63.1 | -0.0017272 | 1.7207 | 0 | 25.8592 | 1.7207 | 2.062 | 2.062 | 2.071 | 1.7263 | 1e-05 | 0.0056012 |  | 2.4897e-10 | -0.0017272 | -63.1 | 1.7263 | 25.8592 | [5 empty] | -0.012235
-63 | -0.0017272 | 1.7207 | 0 | 25.8257 | 1.7207 | 2.062 | 2.062 | 2.0689 | 1.725 | 1e-05 | 0.0043007 |  | 2.5746e-10 | -0.0017272 | -63 | 1.725 | 25.8257 | [5 empty] | -0.015298
-62.9 | -0.0017272 | 1.7207 | 0 | 25.7922 | 1.7207 | 2.062 | 2.062 | 2.0668 | 1.7237 | 1e-05 | 0.0029999 |  | 2.6623e-10 | -0.0017272 | -62.9 | 1.7237 | 25.7922 | [5 empty] | -0.012243
-62.8 | -0.0017272 | 1.7207 | 0 | 25.7587 | 1.7207 | 2.062 | 2.062 | 2.0647 | 1.7224 | 1e-05 | 0.0016987 |  | 2.7529e-10 | -0.0017272 | -62.8 | 1.7224 | 25.7587 | [5 empty] | -0.012243
-62.7 | -0.0017272 | 1.7207 | -0.033692 | 25.7252 | 1.7207 | 2.062 | 2.062 | 2.0626 | 1.7211 | 1e-05 | 0.00039722 |  | 2.8467e-10 | -0.0017272 | -62.7 | 1.7211 | 25.7252 | [5 empty] | -0.012243
-62.6 | -0.0017272 | 1.7174 | -0.033692 | 25.6918 | 1.7174 | 2.0566 | 2.0566 | 2.0606 | 1.7198 | 1e-05 | 1e-05 |  | 2.9437e-10 | -0.0017272 | -62.6 | 1.7198 | 25.6918 | [5 empty] | -0.012252
-62.5 | -0.0017272 | 1.714 | -0.033692 | 25.6583 | 1.714 | 2.0512 | 2.0512 | 2.0585 | 1.7185 | 1e-05 | 1e-05 |  | 3.044e-10 | -0.0017272 | -62.5 | 1.7185 | 25.6583 | [5 empty] | -0.015314
-62.4 | -0.0017272 | 1.7106 | 0 | 25.6248 | 1.7106 | 2.0459 | 2.0459 | 2.0564 | 1.7172 | 1e-05 | 1e-05 |  | 3.1476e-10 | -0.0017272 | -62.4 | 1.7172 | 25.6248 | [5 empty] | -0.018377
-62.3 | -0.0017272 | 1.7106 | 0 | 25.5913 | 1.7106 | 2.0459 | 2.0459 | 2.0543 | 1.7159 | 1e-05 | 1e-05 |  | 3.2549e-10 | -0.0017272 | -62.3 | 1.7159 | 25.5913 | [5 empty] | -0.02144
-62.2 | -0.0017272 | 1.7106 | 0 | 25.5578 | 1.7106 | 2.0459 | 2.0459 | 2.0522 | 1.7146 | 1e-05 | 1e-05 |  | 3.3657e-10 | -0.0017272 | -62.2 | 1.7146 | 25.5578 | [5 empty] | -0.024503
-62.1 | -0.00169 | 1.7106 | -0.033692 | 25.5243 | 1.7106 | 2.0459 | 2.0459 | 2.0502 | 1.7133 | 1e-05 | 0.029642 |  | 3.4804e-10 | -0.00169 | -62.1 | 1.7133 | 25.5243 | [5 empty] | -0.018377
-62 | -0.00169 | 1.7073 | -0.033692 | 25.4908 | 1.7073 | 2.0405 | 2.0405 | 2.0481 | 1.712 | 1e-05 | 0.028339 |  | 3.5989e-10 | -0.00169 | -62 | 1.712 | 25.4908 | [5 empty] | -0.012252
-61.9 | -0.00169 | 1.7039 | -0.033692 | 25.4573 | 1.7039 | 2.0351 | 2.0351 | 2.046 | 1.7107 | 1e-05 | 0.027034 |  | 3.7215e-10 | -0.00169 | -61.9 | 1.7107 | 25.4573 | [5 empty] | -0.0061258
-61.8 | -0.0017272 | 1.7005 | -0.033692 | 25.4238 | 1.7005 | 2.0298 | 2.0298 | 2.0439 | 1.7094 | 1e-05 | 1e-05 |  | 3.8483e-10 | -0.0017272 | -61.8 | 1.7094 | 25.4238 | [5 empty] | -0.0091887
-61.7 | -0.0017272 | 1.6972 | -0.033692 | 25.3903 | 1.6972 | 2.0244 | 2.0244 | 2.0418 | 1.7081 | 1e-05 | 1e-05 |  | 3.9794e-10 | -0.0017272 | -61.7 | 1.7081 | 25.3903 | [5 empty] | -0.0091887
-61.6 | -0.0017272 | 1.6938 | 0.033692 | 25.3568 | 1.6938 | 2.0191 | 2.0191 | 2.0398 | 1.7068 | 1e-05 | 1e-05 |  | 4.1149e-10 | -0.0017272 | -61.6 | 1.7068 | 25.3568 | [5 empty] | -0.0091887
-61.5 | -0.00169 | 1.6972 | 0.033692 | 25.3233 | 1.6972 | 2.0244 | 2.0244 | 2.0377 | 1.7055 | 1e-05 | 0.021815 |  | 4.2551e-10 | -0.00169 | -61.5 | 1.7055 | 25.3233 | [5 empty] | -0.0061258
-61.4 | -0.00169 | 1.7005 | 0.033692 | 25.2898 | 1.7005 | 2.0298 | 2.0298 | 2.0356 | 1.7042 | 1e-05 | 0.020509 |  | 4.4001e-10 | -0.00169 | -61.4 | 1.7042 | 25.2898 | [5 empty] | -0.0030629
-61.3 | -0.00169 | 1.7039 | -0.033692 | 25.2563 | 1.7039 | 2.0351 | 2.0351 | 2.0335 | 1.7029 | 1e-05 | 0.019203 |  | 4.5499e-10 | -0.00169 | -61.3 | 1.7029 | 25.2563 | [5 empty] | -4.3526e-16
-61.2 | -0.00169 | 1.7005 | 0 | 25.2228 | 1.7005 | 2.0298 | 2.0298 | 2.0314 | 1.7016 | 1e-05 | 0.017896 |  | 4.7049e-10 | -0.00169 | -61.2 | 1.7016 | 25.2228 | [5 empty] | 0.0030629
-61.1 | -0.00169 | 1.7005 | 0 | 25.1893 | 1.7005 | 2.0298 | 2.0298 | 2.0294 | 1.7003 | 1e-05 | 0.016589 |  | 4.8652e-10 | -0.00169 | -61.1 | 1.7003 | 25.1893 | [5 empty] | 0.0061258
-61 | -0.0017272 | 1.7005 | 0 | 25.1558 | 1.7005 | 2.0298 | 2.0298 | 2.0273 | 1.699 | 1e-05 | 1e-05 |  | 5.0309e-10 | -0.0017272 | -61 | 1.699 | 25.1558 | [5 empty] | -2.1763e-16
-60.9 | -0.0017272 | 1.7005 | 0 | 25.1223 | 1.7005 | 2.0298 | 2.0298 | 2.0252 | 1.6977 | 1e-05 | 1e-05 |  | 5.2023e-10 | -0.0017272 | -60.9 | 1.6977 | 25.1223 | [5 empty] | -0.0061258
-60.8 | -0.0017272 | 1.7005 | 0 | 25.0888 | 1.7005 | 2.0298 | 2.0298 | 2.0231 | 1.6964 | 1e-05 | 1e-05 |  | 5.3795e-10 | -0.0017272 | -60.8 | 1.6964 | 25.0888 | [5 empty] | -0.012252
-60.7 | -0.00169 | 1.7005 | 0 | 25.0553 | 1.7005 | 2.0298 | 2.0298 | 2.021 | 1.695 | 1e-05 | 0.011359 |  | 5.5628e-10 | -0.00169 | -60.7 | 1.695 | 25.0553 | [5 empty] | -0.0091887
-60.6 | -0.0017272 | 1.7005 | 0 | 25.0218 | 1.7005 | 2.0298 | 2.0298 | 2.019 | 1.6937 | 1e-05 | 1e-05 |  | 5.7522e-10 | -0.0017272 | -60.6 | 1.6937 | 25.0218 | [5 empty] | -0.015314
-60.5 | -0.0017272 | 1.7005 | -0.033692 | 24.9883 | 1.7005 | 2.0298 | 2.0298 | 2.0169 | 1.6924 | 1e-05 | 1e-05 |  | 5.9482e-10 | -0.0017272 | -60.5 | 1.6924 | 24.9883 | [5 empty] | -0.018377
-60.4 | -0.00169 | 1.6972 | -0.033692 | 24.9548 | 1.6972 | 2.0244 | 2.0244 | 2.0148 | 1.6911 | 1e-05 | 0.0074325 |  | 6.1508e-10 | -0.00169 | -60.4 | 1.6911 | 24.9548 | [5 empty] | -0.018377
-60.3 | -0.00169 | 1.6938 | -0.033692 | 24.9213 | 1.6938 | 2.0191 | 2.0191 | 2.0127 | 1.6898 | 1e-05 | 0.006123 |  | 6.3603e-10 | -0.00169 | -60.3 | 1.6898 | 24.9213 | [5 empty] | -0.018377
-60.2 | -0.00169 | 1.6904 | 0 | 24.8878 | 1.6904 | 2.0137 | 2.0137 | 2.0106 | 1.6885 | 1e-05 | 0.0048132 |  | 6.577e-10 | -0.00169 | -60.2 | 1.6885 | 24.8878 | [5 empty] | -0.018377
-60.1 | -0.00169 | 1.6904 | -0.067384 | 24.8543 | 1.6904 | 2.0137 | 2.0137 | 2.0086 | 1.6872 | 1e-05 | 0.003503 |  | 6.801e-10 | -0.00169 | -60.1 | 1.6872 | 24.8543 | [5 empty] | -0.018377
-60 | -0.00169 | 1.6837 | -0.033692 | 24.8208 | 1.6837 | 2.003 | 2.003 | 2.0065 | 1.6859 | 1e-05 | 0.0021925 |  | 7.0327e-10 | -0.00169 | -60 | 1.6859 | 24.8208 | [5 empty] | -0.02144
-59.9 | -0.00169 | 1.6803 | 0 | 24.7873 | 1.6803 | 1.9977 | 1.9977 | 2.0044 | 1.6846 | 1e-05 | 0.00088168 |  | 7.2722e-10 | -0.00169 | -59.9 | 1.6846 | 24.7873 | [5 empty] | -0.02144
-59.8 | -0.00169 | 1.6803 | 0 | 24.7539 | 1.6803 | 1.9977 | 1.9977 | 2.0023 | 1.6833 | 1e-05 | 1e-05 |  | 7.52e-10 | -0.00169 | -59.8 | 1.6833 | 24.7539 | [5 empty] | -0.02144
-59.7 | -0.00169 | 1.6803 | 0 | 24.7204 | 1.6803 | 1.9977 | 1.9977 | 2.0002 | 1.6819 | 1e-05 | 1e-05 |  | 7.7761e-10 | -0.00169 | -59.7 | 1.6819 | 24.7204 | [5 empty] | -0.018377
-59.6 | -0.00169 | 1.6803 | 0 | 24.6869 | 1.6803 | 1.9977 | 1.9977 | 1.9982 | 1.6806 | 1e-05 | 1e-05 |  | 8.041e-10 | -0.00169 | -59.6 | 1.6806 | 24.6869 | [5 empty] | -0.018377
-59.5 | -0.0016528 | 1.6803 | -0.033692 | 24.6534 | 1.6803 | 1.9977 | 1.9977 | 1.9961 | 1.6793 | 1e-05 | 0.032696 |  | 8.3149e-10 | -0.0016528 | -59.5 | 1.6793 | 24.6534 | [5 empty] | -0.012252
-59.4 | -0.00169 | 1.6769 | -0.033692 | 24.6199 | 1.6769 | 1.9923 | 1.9923 | 1.994 | 1.678 | 1e-05 | 1e-05 |  | 8.5982e-10 | -0.00169 | -59.4 | 1.678 | 24.6199 | [5 empty] | -0.012252
-59.3 | -0.00169 | 1.6736 | -0.033692 | 24.5864 | 1.6736 | 1.987 | 1.987 | 1.9919 | 1.6767 | 1e-05 | 1e-05 |  | 8.891e-10 | -0.00169 | -59.3 | 1.6767 | 24.5864 | [5 empty] | -0.012252
-59.2 | -0.00169 | 1.6702 | 0 | 24.5529 | 1.6702 | 1.9817 | 1.9817 | 1.9898 | 1.6754 | 1e-05 | 1e-05 |  | 9.1939e-10 | -0.00169 | -59.2 | 1.6754 | 24.5529 | [5 empty] | -0.012252
-59.1 | -0.00169 | 1.6702 | 0 | 24.5194 | 1.6702 | 1.9817 | 1.9817 | 1.9878 | 1.6741 | 1e-05 | 1e-05 |  | 9.5071e-10 | -0.00169 | -59.1 | 1.6741 | 24.5194 | [5 empty] | -0.015314
-59 | -0.00169 | 1.6702 | 0 | 24.4859 | 1.6702 | 1.9817 | 1.9817 | 1.9857 | 1.6728 | 1e-05 | 1e-05 |  | 9.8309e-10 | -0.00169 | -59 | 1.6728 | 24.4859 | [5 empty] | -0.015314
-58.9 | -0.0016528 | 1.6702 | -0.033692 | 24.4524 | 1.6702 | 1.9817 | 1.9817 | 1.9836 | 1.6714 | 1e-05 | 0.024816 |  | 1.0166e-09 | -0.0016528 | -58.9 | 1.6714 | 24.4524 | [5 empty] | -0.012252
-58.8 | -0.0016528 | 1.6668 | 0 | 24.4189 | 1.6668 | 1.9763 | 1.9763 | 1.9815 | 1.6701 | 1e-05 | 0.023501 |  | 1.0512e-09 | -0.0016528 | -58.8 | 1.6701 | 24.4189 | [5 empty] | -0.0091887
-58.7 | -0.0016528 | 1.6668 | 0 | 24.3854 | 1.6668 | 1.9763 | 1.9763 | 1.9794 | 1.6688 | 1e-05 | 0.022186 |  | 1.087e-09 | -0.0016528 | -58.7 | 1.6688 | 24.3854 | [5 empty] | -0.0061258
-58.6 | -0.00169 | 1.6668 | -0.033692 | 24.3519 | 1.6668 | 1.9763 | 1.9763 | 1.9774 | 1.6675 | 1e-05 | 1e-05 |  | 1.124e-09 | -0.00169 | -58.6 | 1.6675 | 24.3519 | [5 empty] | -0.0091887
-58.5 | -0.00169 | 1.6635 | 0 | 24.3184 | 1.6635 | 1.971 | 1.971 | 1.9753 | 1.6662 | 1e-05 | 1e-05 |  | 1.1623e-09 | -0.00169 | -58.5 | 1.6662 | 24.3184 | [5 empty] | -0.012252
-58.4 | -0.0016528 | 1.6635 | 0 | 24.2849 | 1.6635 | 1.971 | 1.971 | 1.9732 | 1.6649 | 1e-05 | 0.018239 |  | 1.2019e-09 | -0.0016528 | -58.4 | 1.6649 | 24.2849 | [5 empty] | -0.012252
-58.3 | -0.0016528 | 1.6635 | 0 | 24.2514 | 1.6635 | 1.971 | 1.971 | 1.9711 | 1.6635 | 1e-05 | 0.016923 |  | 1.2429e-09 | -0.0016528 | -58.3 | 1.6635 | 24.2514 | [5 empty] | -0.0061258
-58.2 | -0.00169 | 1.6635 | 0 | 24.2179 | 1.6635 | 1.971 | 1.971 | 1.969 | 1.6622 | 1e-05 | 1e-05 |  | 1.2852e-09 | -0.00169 | -58.2 | 1.6622 | 24.2179 | [5 empty] | -0.0091887
-58.1 | -0.00169 | 1.6635 | -0.033692 | 24.1844 | 1.6635 | 1.971 | 1.971 | 1.967 | 1.6609 | 1e-05 | 1e-05 |  | 1.329e-09 | -0.00169 | -58.1 | 1.6609 | 24.1844 | [5 empty] | -0.012252
-58 | -0.0016528 | 1.6601 | -0.033692 | 24.1509 | 1.6601 | 1.9657 | 1.9657 | 1.9649 | 1.6596 | 1e-05 | 0.012972 |  | 1.3743e-09 | -0.0016528 | -58 | 1.6596 | 24.1509 | [5 empty] | -0.0091887
-57.9 | -0.00169 | 1.6567 | 0 | 24.1174 | 1.6567 | 1.9604 | 1.9604 | 1.9628 | 1.6583 | 1e-05 | 1e-05 |  | 1.4211e-09 | -0.00169 | -57.9 | 1.6583 | 24.1174 | [5 empty] | -0.012252
-57.8 | -0.0016528 | 1.6567 | 0.033692 | 24.0839 | 1.6567 | 1.9604 | 1.9604 | 1.9607 | 1.657 | 1e-05 | 0.010336 |  | 1.4695e-09 | -0.0016528 | -57.8 | 1.657 | 24.0839 | [5 empty] | -0.015306
-57.7 | -0.0016528 | 1.6601 | -0.033692 | 24.0504 | 1.6601 | 1.9657 | 1.9657 | 1.9586 | 1.6556 | 1e-05 | 0.0090174 |  | 1.5195e-09 | -0.0016528 | -57.7 | 1.6556 | 24.0504 | [5 empty] | -0.018361
-57.6 | -0.0016528 | 1.6567 | -0.033692 | 24.0169 | 1.6567 | 1.9604 | 1.9604 | 1.9566 | 1.6543 | 1e-05 | 0.0076986 |  | 1.5713e-09 | -0.0016528 | -57.6 | 1.6543 | 24.0169 | [5 empty] | -0.018361
-57.5 | -0.0016528 | 1.6534 | 0 | 23.9834 | 1.6534 | 1.955 | 1.955 | 1.9545 | 1.653 | 1e-05 | 0.0063795 |  | 1.6248e-09 | -0.0016528 | -57.5 | 1.653 | 23.9834 | [5 empty] | -0.018353
-57.4 | -0.0016528 | 1.6534 | -0.033692 | 23.9499 | 1.6534 | 1.955 | 1.955 | 1.9524 | 1.6517 | 1e-05 | 0.00506 |  | 1.6802e-09 | -0.0016528 | -57.4 | 1.6517 | 23.9499 | [5 empty] | -0.018344
-57.3 | -0.0016528 | 1.65 | -0.033601 | 23.9164 | 1.65 | 1.9497 | 1.9497 | 1.9503 | 1.6504 | 1e-05 | 0.0037402 |  | 1.7374e-09 | -0.0016528 | -57.3 | 1.6504 | 23.9164 | [5 empty] | -0.018344
-57.2 | -0.00169 | 1.6466 | -0.033601 | 23.8829 | 1.6466 | 1.9444 | 1.9444 | 1.9482 | 1.649 | 1e-05 | 1e-05 |  | 1.7966e-09 | -0.00169 | -57.2 | 1.649 | 23.8829 | [5 empty] | -0.02447
-57.1 | -0.0016528 | 1.6433 | 0 | 23.8494 | 1.6433 | 1.9392 | 1.9392 | 1.9462 | 1.6477 | 1e-05 | 0.0010994 |  | 1.8578e-09 | -0.0016528 | -57.1 | 1.6477 | 23.8494 | [5 empty] | -0.021407
-57 | -0.0016528 | 1.6433 | -0.033601 | 23.816 | 1.6433 | 1.9392 | 1.9392 | 1.9441 | 1.6464 | 1e-05 | 1e-05 |  | 1.9211e-09 | -0.0016528 | -57 | 1.6464 | 23.816 | [5 empty] | -0.018344
-56.9 | -0.0016528 | 1.6399 | -0.033601 | 23.7825 | 1.6399 | 1.9339 | 1.9339 | 1.942 | 1.6451 | 1e-05 | 1e-05 |  | 1.9865e-09 | -0.0016528 | -56.9 | 1.6451 | 23.7825 | [5 empty] | -0.018344
-56.8 | -0.0016528 | 1.6366 | 0 | 23.749 | 1.6366 | 1.9286 | 1.9286 | 1.9399 | 1.6438 | 1e-05 | 1e-05 |  | 2.0542e-09 | -0.0016528 | -56.8 | 1.6438 | 23.749 | [5 empty] | -0.018336
-56.7 | -0.0016157 | 1.6366 | -0.033692 | 23.7155 | 1.6366 | 1.9286 | 1.9286 | 1.9378 | 1.6424 | 1e-05 | 0.032775 |  | 2.1241e-09 | -0.0016157 | -56.7 | 1.6424 | 23.7155 | [5 empty] | -0.015282
-56.6 | -0.0016157 | 1.6332 | 0 | 23.682 | 1.6332 | 1.9233 | 1.9233 | 1.9358 | 1.6411 | 1e-05 | 0.031453 |  | 2.1965e-09 | -0.0016157 | -56.6 | 1.6411 | 23.682 | [5 empty] | -0.012227
-56.5 | -0.0016528 | 1.6332 | 0 | 23.6485 | 1.6332 | 1.9233 | 1.9233 | 1.9337 | 1.6398 | 1e-05 | 1e-05 |  | 2.2713e-09 | -0.0016528 | -56.5 | 1.6398 | 23.6485 | [5 empty] | -0.015282
-56.4 | -0.0016157 | 1.6332 | 0 | 23.615 | 1.6332 | 1.9233 | 1.9233 | 1.9316 | 1.6385 | 1e-05 | 0.028807 |  | 2.3487e-09 | -0.0016157 | -56.4 | 1.6385 | 23.615 | [5 empty] | -0.012227
-56.3 | -0.0016157 | 1.6332 | -0.033601 | 23.5815 | 1.6332 | 1.9233 | 1.9233 | 1.9295 | 1.6371 | 1e-05 | 0.027484 |  | 2.4287e-09 | -0.0016157 | -56.3 | 1.6371 | 23.5815 | [5 empty] | -0.0091722
-56.2 | -0.0016528 | 1.6298 | 0 | 23.548 | 1.6298 | 1.918 | 1.918 | 1.9274 | 1.6358 | 1e-05 | 1e-05 |  | 2.5114e-09 | -0.0016528 | -56.2 | 1.6358 | 23.548 | [5 empty] | -0.012227
-56.1 | -0.0016528 | 1.6298 | 0 | 23.5145 | 1.6298 | 1.918 | 1.918 | 1.9254 | 1.6345 | 1e-05 | 1e-05 |  | 2.597e-09 | -0.0016528 | -56.1 | 1.6345 | 23.5145 | [5 empty] | -0.009164
-56 | -0.0016528 | 1.6298 | -0.033601 | 23.481 | 1.6298 | 1.918 | 1.918 | 1.9233 | 1.6332 | 1e-05 | 1e-05 |  | 2.6854e-09 | -0.0016528 | -56 | 1.6332 | 23.481 | [5 empty] | -0.012219
-55.9 | -0.0016528 | 1.6265 | 0 | 23.4475 | 1.6265 | 1.9128 | 1.9128 | 1.9212 | 1.6319 | 1e-05 | 1e-05 |  | 2.7769e-09 | -0.0016528 | -55.9 | 1.6319 | 23.4475 | [5 empty] | -0.012219
-55.8 | -0.0016528 | 1.6265 | 0 | 23.414 | 1.6265 | 1.9128 | 1.9128 | 1.9191 | 1.6305 | 1e-05 | 1e-05 |  | 2.8715e-09 | -0.0016528 | -55.8 | 1.6305 | 23.414 | [5 empty] | -0.015273
-55.7 | -0.0016157 | 1.6265 | -0.033601 | 23.3805 | 1.6265 | 1.9128 | 1.9128 | 1.917 | 1.6292 | 1e-05 | 0.019536 |  | 2.9693e-09 | -0.0016157 | -55.7 | 1.6292 | 23.3805 | [5 empty] | -0.012219
-55.6 | -0.0016157 | 1.6231 | 0 | 23.347 | 1.6231 | 1.9075 | 1.9075 | 1.915 | 1.6279 | 1e-05 | 0.01821 |  | 3.0705e-09 | -0.0016157 | -55.6 | 1.6279 | 23.347 | [5 empty] | -0.012219
-55.5 | -0.0016157 | 1.6231 | -0.033601 | 23.3135 | 1.6231 | 1.9075 | 1.9075 | 1.9129 | 1.6265 | 1e-05 | 0.016884 |  | 3.175e-09 | -0.0016157 | -55.5 | 1.6265 | 23.3135 | [5 empty] | -0.012219
-55.4 | -0.0016157 | 1.6197 | 0 | 23.28 | 1.6197 | 1.9022 | 1.9022 | 1.9108 | 1.6252 | 1e-05 | 0.015557 |  | 3.2832e-09 | -0.0016157 | -55.4 | 1.6252 | 23.28 | [5 empty] | -0.009164
-55.3 | -0.0016157 | 1.6197 | -0.033601 | 23.2465 | 1.6197 | 1.9022 | 1.9022 | 1.9087 | 1.6239 | 1e-05 | 0.01423 |  | 3.395e-09 | -0.0016157 | -55.3 | 1.6239 | 23.2465 | [5 empty] | -0.009164
-55.2 | -0.0016157 | 1.6164 | 0 | 23.213 | 1.6164 | 1.897 | 1.897 | 1.9066 | 1.6226 | 1e-05 | 0.012903 |  | 3.5107e-09 | -0.0016157 | -55.2 | 1.6226 | 23.213 | [5 empty] | -0.009164
-55.1 | -0.0016157 | 1.6164 | 0 | 23.1795 | 1.6164 | 1.897 | 1.897 | 1.9046 | 1.6212 | 1e-05 | 0.011575 |  | 3.6303e-09 | -0.0016157 | -55.1 | 1.6212 | 23.1795 | [5 empty] | -0.0061093
-55 | -0.0016528 | 1.6164 | 0 | 23.146 | 1.6164 | 1.897 | 1.897 | 1.9025 | 1.6199 | 1e-05 | 1e-05 |  | 3.7539e-09 | -0.0016528 | -55 | 1.6199 | 23.146 | [5 empty] | -0.009164
-54.9 | -0.0016157 | 1.6164 | 0 | 23.1125 | 1.6164 | 1.897 | 1.897 | 1.9004 | 1.6186 | 1e-05 | 0.008919 |  | 3.8818e-09 | -0.0016157 | -54.9 | 1.6186 | 23.1125 | [5 empty] | -0.0061093
-54.8 | -0.0016528 | 1.6164 | 0 | 23.079 | 1.6164 | 1.897 | 1.897 | 1.8983 | 1.6173 | 1e-05 | 1e-05 |  | 4.014e-09 | -0.0016528 | -54.8 | 1.6173 | 23.079 | [5 empty] | -0.009164
-54.7 | -0.0016157 | 1.6164 | 0 | 23.0455 | 1.6164 | 1.897 | 1.897 | 1.8962 | 1.6159 | 1e-05 | 0.0062611 |  | 4.1508e-09 | -0.0016157 | -54.7 | 1.6159 | 23.0455 | [5 empty] | -0.0061093
-54.6 | -0.0016157 | 1.6164 | 0 | 23.012 | 1.6164 | 1.897 | 1.897 | 1.8942 | 1.6146 | 1e-05 | 0.0049316 |  | 4.2922e-09 | -0.0016157 | -54.6 | 1.6146 | 23.012 | [5 empty] | -0.0061093
-54.5 | -0.0016157 | 1.6164 | -0.033601 | 22.9785 | 1.6164 | 1.897 | 1.897 | 1.8921 | 1.6133 | 1e-05 | 0.0036017 |  | 4.4384e-09 | -0.0016157 | -54.5 | 1.6133 | 22.9785 | [5 empty] | -0.0061093
-54.4 | -0.0016157 | 1.613 | 0 | 22.945 | 1.613 | 1.8917 | 1.8917 | 1.89 | 1.6119 | 1e-05 | 0.0022715 |  | 4.5896e-09 | -0.0016157 | -54.4 | 1.6119 | 22.945 | [5 empty] | -0.0091722
-54.3 | -0.0016157 | 1.613 | -0.033601 | 22.9115 | 1.613 | 1.8917 | 1.8917 | 1.8879 | 1.6106 | 1e-05 | 0.00094088 |  | 4.7459e-09 | -0.0016157 | -54.3 | 1.6106 | 22.9115 | [5 empty] | -0.012235
-54.2 | -0.0016157 | 1.6097 | 0 | 22.8781 | 1.6097 | 1.8865 | 1.8865 | 1.8858 | 1.6093 | 1e-05 | 1e-05 |  | 4.9076e-09 | -0.0016157 | -54.2 | 1.6093 | 22.8781 | [5 empty] | -0.012235
-54.1 | -0.0016157 | 1.6097 | 0 | 22.8446 | 1.6097 | 1.8865 | 1.8865 | 1.8838 | 1.6079 | 1e-05 | 1e-05 |  | 5.0747e-09 | -0.0016157 | -54.1 | 1.6079 | 22.8446 | [5 empty] | -0.012235
-54 | -0.0016157 | 1.6097 | 0 | 22.8111 | 1.6097 | 1.8865 | 1.8865 | 1.8817 | 1.6066 | 1e-05 | 1e-05 |  | 5.2476e-09 | -0.0016157 | -54 | 1.6066 | 22.8111 | [5 empty] | -0.015298
-53.9 | -0.0016157 | 1.6097 | -0.033692 | 22.7776 | 1.6097 | 1.8865 | 1.8865 | 1.8796 | 1.6053 | 1e-05 | 1e-05 |  | 5.4263e-09 | -0.0016157 | -53.9 | 1.6053 | 22.7776 | [5 empty] | -0.015306
-53.8 | -0.0016157 | 1.6063 | -0.033692 | 22.7441 | 1.6063 | 1.8812 | 1.8812 | 1.8775 | 1.6039 | 1e-05 | 1e-05 |  | 5.6112e-09 | -0.0016157 | -53.8 | 1.6039 | 22.7441 | [5 empty] | -0.015306
-53.7 | -0.0016157 | 1.6029 | 0 | 22.7106 | 1.6029 | 1.8759 | 1.8759 | 1.8754 | 1.6026 | 1e-05 | 1e-05 |  | 5.8023e-09 | -0.0016157 | -53.7 | 1.6026 | 22.7106 | [5 empty] | -0.012252
-53.6 | -0.0016157 | 1.6029 | 0 | 22.6771 | 1.6029 | 1.8759 | 1.8759 | 1.8734 | 1.6013 | 1e-05 | 1e-05 |  | 6e-09 | -0.0016157 | -53.6 | 1.6013 | 22.6771 | [5 empty] | -0.015314
-53.5 | -0.0016157 | 1.6029 | -0.033692 | 22.6436 | 1.6029 | 1.8759 | 1.8759 | 1.8713 | 1.5999 | 1e-05 | 1e-05 |  | 6.2044e-09 | -0.0016157 | -53.5 | 1.5999 | 22.6436 | [5 empty] | -0.015314
-53.4 | -0.0016157 | 1.5996 | -0.033692 | 22.6101 | 1.5996 | 1.8707 | 1.8707 | 1.8692 | 1.5986 | 1e-05 | 1e-05 |  | 6.4157e-09 | -0.0016157 | -53.4 | 1.5986 | 22.6101 | [5 empty] | -0.018377
-53.3 | -0.0015785 | 1.5962 | 0 | 22.5766 | 1.5962 | 1.8654 | 1.8654 | 1.8671 | 1.5973 | 1e-05 | 0.024676 |  | 6.6342e-09 | -0.0015785 | -53.3 | 1.5973 | 22.5766 | [5 empty] | -0.012252
-53.2 | -0.0015785 | 1.5962 | 0 | 22.5431 | 1.5962 | 1.8654 | 1.8654 | 1.865 | 1.5959 | 1e-05 | 0.023341 |  | 6.8602e-09 | -0.0015785 | -53.2 | 1.5959 | 22.5431 | [5 empty] | -0.0061258
-53.1 | -0.0016157 | 1.5962 | -0.033692 | 22.5096 | 1.5962 | 1.8654 | 1.8654 | 1.863 | 1.5946 | 1e-05 | 1e-05 |  | 7.0939e-09 | -0.0016157 | -53.1 | 1.5946 | 22.5096 | [5 empty] | -0.0061258
-53 | -0.0016157 | 1.5928 | 0 | 22.4761 | 1.5928 | 1.8602 | 1.8602 | 1.8609 | 1.5933 | 1e-05 | 1e-05 |  | 7.3356e-09 | -0.0016157 | -53 | 1.5933 | 22.4761 | [5 empty] | -0.0061258
-52.9 | -0.0015785 | 1.5928 | -0.033692 | 22.4426 | 1.5928 | 1.8602 | 1.8602 | 1.8588 | 1.5919 | 1e-05 | 0.019335 |  | 7.5854e-09 | -0.0015785 | -52.9 | 1.5919 | 22.4426 | [5 empty] | -0.0030629
-52.8 | -0.0015785 | 1.5894 | 0.033692 | 22.4091 | 1.5894 | 1.8549 | 1.8549 | 1.8567 | 1.5906 | 1e-05 | 0.017999 |  | 7.8438e-09 | -0.0015785 | -52.8 | 1.5906 | 22.4091 | [5 empty] | 2.1763e-16
-52.7 | -0.0016157 | 1.5928 | 0.033692 | 22.3756 | 1.5928 | 1.8602 | 1.8602 | 1.8546 | 1.5893 | 1e-05 | 1e-05 |  | 8.111e-09 | -0.0016157 | -52.7 | 1.5893 | 22.3756 | [5 empty] | -0.0030629
-52.6 | -0.0016157 | 1.5962 | 0 | 22.3421 | 1.5962 | 1.8654 | 1.8654 | 1.8526 | 1.5879 | 1e-05 | 1e-05 |  | 8.3873e-09 | -0.0016157 | -52.6 | 1.5879 | 22.3421 | [5 empty] | -0.0061258
-52.5 | -0.0015785 | 1.5962 | 0 | 22.3086 | 1.5962 | 1.8654 | 1.8654 | 1.8505 | 1.5866 | 1e-05 | 0.013988 |  | 8.673e-09 | -0.0015785 | -52.5 | 1.5866 | 22.3086 | [5 empty] | 2.1763e-16
-52.4 | -0.0016157 | 1.5962 | 0 | 22.2751 | 1.5962 | 1.8654 | 1.8654 | 1.8484 | 1.5853 | 1e-05 | 1e-05 |  | 8.9684e-09 | -0.0016157 | -52.4 | 1.5853 | 22.2751 | [5 empty] | -0.0030629
-52.3 | -0.0015785 | 1.5962 | 0 | 22.2416 | 1.5962 | 1.8654 | 1.8654 | 1.8463 | 1.5839 | 1e-05 | 0.011312 |  | 9.2739e-09 | -0.0015785 | -52.3 | 1.5839 | 22.2416 | [5 empty] | 2.1763e-16
-52.2 | -0.0016157 | 1.5962 | -0.033692 | 22.2081 | 1.5962 | 1.8654 | 1.8654 | 1.8442 | 1.5826 | 1e-05 | 1e-05 |  | 9.5898e-09 | -0.0016157 | -52.2 | 1.5826 | 22.2081 | [5 empty] | -0.0061258
-52.1 | -0.0016157 | 1.5928 | -0.033692 | 22.1746 | 1.5928 | 1.8602 | 1.8602 | 1.8422 | 1.5812 | 1e-05 | 1e-05 |  | 9.9165e-09 | -0.0016157 | -52.1 | 1.5812 | 22.1746 | [5 empty] | -0.012252
-52 | -0.0016157 | 1.5894 | 0.033692 | 22.1411 | 1.5894 | 1.8549 | 1.8549 | 1.8401 | 1.5799 | 1e-05 | 1e-05 |  | 1.0254e-08 | -0.0016157 | -52 | 1.5799 | 22.1411 | [5 empty] | -0.015314
-51.9 | -0.0016157 | 1.5928 | -0.033692 | 22.1076 | 1.5928 | 1.8602 | 1.8602 | 1.838 | 1.5786 | 1e-05 | 1e-05 |  | 1.0604e-08 | -0.0016157 | -51.9 | 1.5786 | 22.1076 | [5 empty] | -0.018377
-51.8 | -0.0015785 | 1.5894 | 0 | 22.0741 | 1.5894 | 1.8549 | 1.8549 | 1.8359 | 1.5772 | 1e-05 | 0.0046165 |  | 1.0965e-08 | -0.0015785 | -51.8 | 1.5772 | 22.0741 | [5 empty] | -0.018377
-51.7 | -0.0015785 | 1.5894 | -0.033692 | 22.0406 | 1.5894 | 1.8549 | 1.8549 | 1.8339 | 1.5759 | 1e-05 | 0.0032762 |  | 1.1338e-08 | -0.0015785 | -51.7 | 1.5759 | 22.0406 | [5 empty] | -0.018377
-51.6 | -0.0015785 | 1.5861 | -0.033692 | 22.0071 | 1.5861 | 1.8497 | 1.8497 | 1.8318 | 1.5745 | 1e-05 | 0.0019356 |  | 1.1725e-08 | -0.0015785 | -51.6 | 1.5745 | 22.0071 | [5 empty] | -0.015314
-51.5 | -0.0015785 | 1.5827 | -0.033692 | 21.9736 | 1.5827 | 1.8445 | 1.8445 | 1.8297 | 1.5732 | 1e-05 | 0.00059452 |  | 1.2124e-08 | -0.0015785 | -51.5 | 1.5732 | 21.9736 | [5 empty] | -0.012252
-51.4 | -0.0016157 | 1.5793 | -0.033692 | 21.9402 | 1.5793 | 1.8392 | 1.8392 | 1.8276 | 1.5719 | 1e-05 | 1e-05 |  | 1.2537e-08 | -0.0016157 | -51.4 | 1.5719 | 21.9402 | [5 empty] | -0.018377
-51.3 | -0.0015785 | 1.576 | 0 | 21.9067 | 1.576 | 1.834 | 1.834 | 1.8255 | 1.5705 | 1e-05 | 1e-05 |  | 1.2964e-08 | -0.0015785 | -51.3 | 1.5705 | 21.9067 | [5 empty] | -0.015314
-51.2 | -0.0015785 | 1.576 | 0 | 21.8732 | 1.576 | 1.834 | 1.834 | 1.8235 | 1.5692 | 1e-05 | 1e-05 |  | 1.3406e-08 | -0.0015785 | -51.2 | 1.5692 | 21.8732 | [5 empty] | -0.015314
-51.1 | -0.0015785 | 1.576 | 0 | 21.8397 | 1.576 | 1.834 | 1.834 | 1.8214 | 1.5678 | 1e-05 | 1e-05 |  | 1.3862e-08 | -0.0015785 | -51.1 | 1.5678 | 21.8397 | [5 empty] | -0.015314
-51 | -0.0015785 | 1.576 | 0 | 21.8062 | 1.576 | 1.834 | 1.834 | 1.8193 | 1.5665 | 1e-05 | 1e-05 |  | 1.4334e-08 | -0.0015785 | -51 | 1.5665 | 21.8062 | [5 empty] | -0.012252
-50.9 | -0.0015785 | 1.576 | -0.033692 | 21.7727 | 1.576 | 1.834 | 1.834 | 1.8172 | 1.5651 | 1e-05 | 1e-05 |  | 1.4823e-08 | -0.0015785 | -50.9 | 1.5651 | 21.7727 | [5 empty] | -0.0091887
-50.8 | -0.0015785 | 1.5726 | 0 | 21.7392 | 1.5726 | 1.8288 | 1.8288 | 1.8151 | 1.5638 | 1e-05 | 1e-05 |  | 1.5328e-08 | -0.0015785 | -50.8 | 1.5638 | 21.7392 | [5 empty] | -0.0061258
-50.7 | -0.0015785 | 1.5726 | 0 | 21.7057 | 1.5726 | 1.8288 | 1.8288 | 1.8131 | 1.5625 | 1e-05 | 1e-05 |  | 1.585e-08 | -0.0015785 | -50.7 | 1.5625 | 21.7057 | [5 empty] | -0.0091887
-50.6 | -0.0015785 | 1.5726 | -0.033692 | 21.6722 | 1.5726 | 1.8288 | 1.8288 | 1.811 | 1.5611 | 1e-05 | 1e-05 |  | 1.639e-08 | -0.0015785 | -50.6 | 1.5611 | 21.6722 | [5 empty] | -0.0091887
-50.5 | -0.0015785 | 1.5692 | 0 | 21.6387 | 1.5692 | 1.8235 | 1.8235 | 1.8089 | 1.5598 | 1e-05 | 1e-05 |  | 1.6948e-08 | -0.0015785 | -50.5 | 1.5598 | 21.6387 | [5 empty] | -0.0091887
-50.4 | -0.0015785 | 1.5692 | 0 | 21.6052 | 1.5692 | 1.8235 | 1.8235 | 1.8068 | 1.5584 | 1e-05 | 1e-05 |  | 1.7525e-08 | -0.0015785 | -50.4 | 1.5584 | 21.6052 | [5 empty] | -0.012252
-50.3 | -0.0015785 | 1.5692 | 0 | 21.5717 | 1.5692 | 1.8235 | 1.8235 | 1.8047 | 1.5571 | 1e-05 | 1e-05 |  | 1.8122e-08 | -0.0015785 | -50.3 | 1.5571 | 21.5717 | [5 empty] | -0.012252
-50.2 | -0.0015785 | 1.5692 | -0.033692 | 21.5382 | 1.5692 | 1.8235 | 1.8235 | 1.8027 | 1.5557 | 1e-05 | 1e-05 |  | 1.8739e-08 | -0.0015785 | -50.2 | 1.5557 | 21.5382 | [5 empty] | -0.015314
-50.1 | -0.0015785 | 1.5659 | 0 | 21.5047 | 1.5659 | 1.8183 | 1.8183 | 1.8006 | 1.5544 | 1e-05 | 1e-05 |  | 1.9378e-08 | -0.0015785 | -50.1 | 1.5544 | 21.5047 | [5 empty] | -0.018377
-50 | -0.0015413 | 1.5659 | 0 | 21.4712 | 1.5659 | 1.8183 | 1.8183 | 1.7985 | 1.553 | 1e-05 | 0.017494 |  | 2.0038e-08 | -0.0015413 | -50 | 1.553 | 21.4712 | [5 empty] | -0.015314
-49.9 | -0.0015785 | 1.5659 | -0.033692 | 21.4377 | 1.5659 | 1.8183 | 1.8183 | 1.7964 | 1.5517 | 1e-05 | 1e-05 |  | 2.072e-08 | -0.0015785 | -49.9 | 1.5517 | 21.4377 | [5 empty] | -0.018377
-49.8 | -0.0015785 | 1.5625 | -0.033692 | 21.4042 | 1.5625 | 1.8131 | 1.8131 | 1.7943 | 1.5503 | 1e-05 | 1e-05 |  | 2.1426e-08 | -0.0015785 | -49.8 | 1.5503 | 21.4042 | [5 empty] | -0.02144
-49.7 | -0.0015785 | 1.5591 | -0.033692 | 21.3707 | 1.5591 | 1.8079 | 1.8079 | 1.7923 | 1.549 | 1e-05 | 1e-05 |  | 2.2156e-08 | -0.0015785 | -49.7 | 1.549 | 21.3707 | [5 empty] | -0.024503
-49.6 | -0.0015413 | 1.5558 | -0.033692 | 21.3372 | 1.5558 | 1.8027 | 1.8027 | 1.7902 | 1.5476 | 1e-05 | 0.012102 |  | 2.2911e-08 | -0.0015413 | -49.6 | 1.5476 | 21.3372 | [5 empty] | -0.02144
-49.5 | -0.0015785 | 1.5524 | 0 | 21.3037 | 1.5524 | 1.7975 | 1.7975 | 1.7881 | 1.5463 | 1e-05 | 1e-05 |  | 2.3691e-08 | -0.0015785 | -49.5 | 1.5463 | 21.3037 | [5 empty] | -0.024503
-49.4 | -0.0015785 | 1.5524 | -0.033692 | 21.2702 | 1.5524 | 1.7975 | 1.7975 | 1.786 | 1.5449 | 1e-05 | 1e-05 |  | 2.4498e-08 | -0.0015785 | -49.4 | 1.5449 | 21.2702 | [5 empty] | -0.027566
-49.3 | -0.0015413 | 1.549 | -0.033692 | 21.2367 | 1.549 | 1.7923 | 1.7923 | 1.7839 | 1.5436 | 1e-05 | 0.0080549 |  | 2.5333e-08 | -0.0015413 | -49.3 | 1.5436 | 21.2367 | [5 empty] | -0.024503
-49.2 | -0.0015413 | 1.5456 | -0.033692 | 21.2032 | 1.5456 | 1.7871 | 1.7871 | 1.7819 | 1.5422 | 1e-05 | 0.006705 |  | 2.6196e-08 | -0.0015413 | -49.2 | 1.5422 | 21.2032 | [5 empty] | -0.02144
-49.1 | -0.0015413 | 1.5423 | 0 | 21.1697 | 1.5423 | 1.7819 | 1.7819 | 1.7798 | 1.5409 | 1e-05 | 0.0053546 |  | 2.7088e-08 | -0.0015413 | -49.1 | 1.5409 | 21.1697 | [5 empty] | -0.018377
-49 | -0.0015413 | 1.5423 | -0.033692 | 21.1362 | 1.5423 | 1.7819 | 1.7819 | 1.7777 | 1.5395 | 1e-05 | 0.0040039 |  | 2.8011e-08 | -0.0015413 | -49 | 1.5395 | 21.1362 | [5 empty] | -0.015314
-48.9 | -0.0015413 | 1.5389 | -0.033692 | 21.1027 | 1.5389 | 1.7767 | 1.7767 | 1.7756 | 1.5382 | 1e-05 | 0.0026527 |  | 2.8965e-08 | -0.0015413 | -48.9 | 1.5382 | 21.1027 | [5 empty] | -0.015314
-48.8 | -0.0015413 | 1.5355 | 0 | 21.0692 | 1.5355 | 1.7715 | 1.7715 | 1.7735 | 1.5368 | 1e-05 | 0.0013012 |  | 2.9952e-08 | -0.0015413 | -48.8 | 1.5368 | 21.0692 | [5 empty] | -0.012252
-48.7 | -0.0015413 | 1.5355 | 0 | 21.0357 | 1.5355 | 1.7715 | 1.7715 | 1.7715 | 1.5355 | 1e-05 | 1e-05 |  | 3.0972e-08 | -0.0015413 | -48.7 | 1.5355 | 21.0357 | [5 empty] | -0.0091887
-48.6 | -0.0015413 | 1.5355 | 0 | 21.0023 | 1.5355 | 1.7715 | 1.7715 | 1.7694 | 1.5341 | 1e-05 | 1e-05 |  | 3.2027e-08 | -0.0015413 | -48.6 | 1.5341 | 21.0023 | [5 empty] | -0.0061258
-48.5 | -0.0015413 | 1.5355 | 0 | 20.9688 | 1.5355 | 1.7715 | 1.7715 | 1.7673 | 1.5328 | 1e-05 | 1e-05 |  | 3.3118e-08 | -0.0015413 | -48.5 | 1.5328 | 20.9688 | [5 empty] | -0.0091887
-48.4 | -0.0015413 | 1.5355 | 0 | 20.9353 | 1.5355 | 1.7715 | 1.7715 | 1.7652 | 1.5314 | 1e-05 | 1e-05 |  | 3.4246e-08 | -0.0015413 | -48.4 | 1.5314 | 20.9353 | [5 empty] | -0.0061258
-48.3 | -0.0015413 | 1.5355 | 0 | 20.9018 | 1.5355 | 1.7715 | 1.7715 | 1.7631 | 1.5301 | 1e-05 | 1e-05 |  | 3.5412e-08 | -0.0015413 | -48.3 | 1.5301 | 20.9018 | [5 empty] | -0.0061258
-48.2 | -0.0015413 | 1.5355 | 0 | 20.8683 | 1.5355 | 1.7715 | 1.7715 | 1.7611 | 1.5287 | 1e-05 | 1e-05 |  | 3.6619e-08 | -0.0015413 | -48.2 | 1.5287 | 20.8683 | [5 empty] | -0.0061258
-48.1 | -0.0015413 | 1.5355 | 0 | 20.8348 | 1.5355 | 1.7715 | 1.7715 | 1.759 | 1.5274 | 1e-05 | 1e-05 |  | 3.7866e-08 | -0.0015413 | -48.1 | 1.5274 | 20.8348 | [5 empty] | -0.0061258
-48 | -0.0015413 | 1.5355 | -0.033692 | 20.8013 | 1.5355 | 1.7715 | 1.7715 | 1.7569 | 1.526 | 1e-05 | 1e-05 |  | 3.9156e-08 | -0.0015413 | -48 | 1.526 | 20.8013 | [5 empty] | -0.0061258
-47.9 | -0.0015413 | 1.5322 | 0 | 20.7678 | 1.5322 | 1.7664 | 1.7664 | 1.7548 | 1.5247 | 1e-05 | 1e-05 |  | 4.049e-08 | -0.0015413 | -47.9 | 1.5247 | 20.7678 | [5 empty] | -0.0091887
-47.8 | -0.0015413 | 1.5322 | -0.033692 | 20.7343 | 1.5322 | 1.7664 | 1.7664 | 1.7527 | 1.5233 | 1e-05 | 1e-05 |  | 4.1869e-08 | -0.0015413 | -47.8 | 1.5233 | 20.7343 | [5 empty] | -0.012252
-47.7 | -0.0015413 | 1.5288 | 0 | 20.7008 | 1.5288 | 1.7612 | 1.7612 | 1.7507 | 1.5219 | 1e-05 | 1e-05 |  | 4.3295e-08 | -0.0015413 | -47.7 | 1.5219 | 20.7008 | [5 empty] | -0.015314
-47.6 | -0.0015413 | 1.5288 | 0 | 20.6673 | 1.5288 | 1.7612 | 1.7612 | 1.7486 | 1.5206 | 1e-05 | 1e-05 |  | 4.477e-08 | -0.0015413 | -47.6 | 1.5206 | 20.6673 | [5 empty] | -0.018377
-47.5 | -0.0015413 | 1.5288 | 0 | 20.6338 | 1.5288 | 1.7612 | 1.7612 | 1.7465 | 1.5192 | 1e-05 | 1e-05 |  | 4.6295e-08 | -0.0015413 | -47.5 | 1.5192 | 20.6338 | [5 empty] | -0.02144
-47.4 | -0.0015041 | 1.5288 | -0.033692 | 20.6003 | 1.5288 | 1.7612 | 1.7612 | 1.7444 | 1.5179 | 1e-05 | 0.019399 |  | 4.7872e-08 | -0.0015041 | -47.4 | 1.5179 | 20.6003 | [5 empty] | -0.018377
-47.3 | -0.0015413 | 1.5254 | -0.033692 | 20.5668 | 1.5254 | 1.756 | 1.756 | 1.7423 | 1.5165 | 1e-05 | 1e-05 |  | 4.9503e-08 | -0.0015413 | -47.3 | 1.5165 | 20.5668 | [5 empty] | -0.02144
-47.2 | -0.0015041 | 1.5221 | -0.033692 | 20.5333 | 1.5221 | 1.7508 | 1.7508 | 1.7403 | 1.5152 | 1e-05 | 0.016683 |  | 5.1189e-08 | -0.0015041 | -47.2 | 1.5152 | 20.5333 | [5 empty] | -0.018377
-47.1 | -0.0015413 | 1.5187 | -0.033692 | 20.4998 | 1.5187 | 1.7457 | 1.7457 | 1.7382 | 1.5138 | 1e-05 | 1e-05 |  | 5.2933e-08 | -0.0015413 | -47.1 | 1.5138 | 20.4998 | [5 empty] | -0.02144
-47 | -0.0015413 | 1.5153 | -0.033692 | 20.4663 | 1.5153 | 1.7405 | 1.7405 | 1.7361 | 1.5124 | 1e-05 | 1e-05 |  | 5.4736e-08 | -0.0015413 | -47 | 1.5124 | 20.4663 | [5 empty] | -0.02144
-46.9 | -0.0015413 | 1.512 | 0 | 20.4328 | 1.512 | 1.7353 | 1.7353 | 1.734 | 1.5111 | 1e-05 | 1e-05 |  | 5.66e-08 | -0.0015413 | -46.9 | 1.5111 | 20.4328 | [5 empty] | -0.024503
-46.8 | -0.0015041 | 1.512 | -0.033692 | 20.3993 | 1.512 | 1.7353 | 1.7353 | 1.7319 | 1.5097 | 1e-05 | 0.011247 |  | 5.8528e-08 | -0.0015041 | -46.8 | 1.5097 | 20.3993 | [5 empty] | -0.02144
-46.7 | -0.0015041 | 1.5086 | 0 | 20.3658 | 1.5086 | 1.7302 | 1.7302 | 1.7299 | 1.5084 | 1e-05 | 0.0098874 |  | 6.0522e-08 | -0.0015041 | -46.7 | 1.5084 | 20.3658 | [5 empty] | -0.018377
-46.6 | -0.0015041 | 1.5086 | -0.033692 | 20.3323 | 1.5086 | 1.7302 | 1.7302 | 1.7278 | 1.507 | 1e-05 | 0.008527 |  | 6.2584e-08 | -0.0015041 | -46.6 | 1.507 | 20.3323 | [5 empty] | -0.015314
-46.5 | -0.0015041 | 1.5052 | 0 | 20.2988 | 1.5052 | 1.725 | 1.725 | 1.7257 | 1.5056 | 1e-05 | 0.0071662 |  | 6.4715e-08 | -0.0015041 | -46.5 | 1.5056 | 20.2988 | [5 empty] | -0.012252
-46.4 | -0.0015041 | 1.5052 | -0.033692 | 20.2653 | 1.5052 | 1.725 | 1.725 | 1.7236 | 1.5043 | 1e-05 | 0.005805 |  | 6.692e-08 | -0.0015041 | -46.4 | 1.5043 | 20.2653 | [5 empty] | -0.0091887
-46.3 | -0.0015041 | 1.5018 | 0 | 20.2318 | 1.5018 | 1.7199 | 1.7199 | 1.7215 | 1.5029 | 1e-05 | 0.0044434 |  | 6.9199e-08 | -0.0015041 | -46.3 | 1.5029 | 20.2318 | [5 empty] | -0.0091887
-46.2 | -0.0015041 | 1.5018 | 0 | 20.1983 | 1.5018 | 1.7199 | 1.7199 | 1.7195 | 1.5016 | 1e-05 | 0.0030813 |  | 7.1557e-08 | -0.0015041 | -46.2 | 1.5016 | 20.1983 | [5 empty] | -0.0061258
-46.1 | -0.0015041 | 1.5018 | 0 | 20.1648 | 1.5018 | 1.7199 | 1.7199 | 1.7174 | 1.5002 | 1e-05 | 0.0017189 |  | 7.3994e-08 | -0.0015041 | -46.1 | 1.5002 | 20.1648 | [5 empty] | -0.0061258
-46 | -0.0015041 | 1.5018 | 0 | 20.1313 | 1.5018 | 1.7199 | 1.7199 | 1.7153 | 1.4988 | 1e-05 | 0.00035603 |  | 7.6515e-08 | -0.0015041 | -46 | 1.4988 | 20.1313 | [5 empty] | -0.0030629
-45.9 | -0.0015413 | 1.5018 | 0 | 20.0978 | 1.5018 | 1.7199 | 1.7199 | 1.7132 | 1.4975 | 1e-05 | 1e-05 |  | 7.9121e-08 | -0.0015413 | -45.9 | 1.4975 | 20.0978 | [5 empty] | -0.0061258
-45.8 | -0.0015041 | 1.5018 | 0 | 20.0644 | 1.5018 | 1.7199 | 1.7199 | 1.7111 | 1.4961 | 1e-05 | 1e-05 |  | 8.1816e-08 | -0.0015041 | -45.8 | 1.4961 | 20.0644 | [5 empty] | -0.0030629
-45.7 | -0.0015041 | 1.5018 | 0 | 20.0309 | 1.5018 | 1.7199 | 1.7199 | 1.7091 | 1.4947 | 1e-05 | 1e-05 |  | 8.4603e-08 | -0.0015041 | -45.7 | 1.4947 | 20.0309 | [5 empty] | -0.0061176
-45.6 | -0.0015041 | 1.5018 | 0 | 19.9974 | 1.5018 | 1.7199 | 1.7199 | 1.707 | 1.4934 | 1e-05 | 1e-05 |  | 8.7485e-08 | -0.0015041 | -45.6 | 1.4934 | 19.9974 | [5 empty] | -0.0061176
-45.5 | -0.0015041 | 1.5018 | 0 | 19.9639 | 1.5018 | 1.7199 | 1.7199 | 1.7049 | 1.492 | 1e-05 | 1e-05 |  | 9.0465e-08 | -0.0015041 | -45.5 | 1.492 | 19.9639 | [5 empty] | -0.0091722
-45.4 | -0.0015041 | 1.5018 | -0.033692 | 19.9304 | 1.5018 | 1.7199 | 1.7199 | 1.7028 | 1.4907 | 1e-05 | 1e-05 |  | 9.3547e-08 | -0.0015041 | -45.4 | 1.4907 | 19.9304 | [5 empty] | -0.0091722
-45.3 | -0.0015041 | 1.4985 | 0 | 19.8969 | 1.4985 | 1.7148 | 1.7148 | 1.7007 | 1.4893 | 1e-05 | 1e-05 |  | 9.6733e-08 | -0.0015041 | -45.3 | 1.4893 | 19.8969 | [5 empty] | -0.0091722
-45.2 | -0.0015041 | 1.4985 | -0.033601 | 19.8634 | 1.4985 | 1.7148 | 1.7148 | 1.6987 | 1.4879 | 1e-05 | 1e-05 |  | 1.0003e-07 | -0.0015041 | -45.2 | 1.4879 | 19.8634 | [5 empty] | -0.012227
-45.1 | -0.0015041 | 1.4951 | 0 | 19.8299 | 1.4951 | 1.7096 | 1.7096 | 1.6966 | 1.4866 | 1e-05 | 1e-05 |  | 1.0344e-07 | -0.0015041 | -45.1 | 1.4866 | 19.8299 | [5 empty] | -0.012227
-45 | -0.0015041 | 1.4951 | -0.033601 | 19.7964 | 1.4951 | 1.7096 | 1.7096 | 1.6945 | 1.4852 | 1e-05 | 1e-05 |  | 1.0696e-07 | -0.0015041 | -45 | 1.4852 | 19.7964 | [5 empty] | -0.012227
-44.9 | -0.0015041 | 1.4918 | 0 | 19.7629 | 1.4918 | 1.7045 | 1.7045 | 1.6924 | 1.4838 | 1e-05 | 1e-05 |  | 1.106e-07 | -0.0015041 | -44.9 | 1.4838 | 19.7629 | [5 empty] | -0.015282
-44.8 | -0.0015041 | 1.4918 | 0 | 19.7294 | 1.4918 | 1.7045 | 1.7045 | 1.6903 | 1.4824 | 1e-05 | 1e-05 |  | 1.1437e-07 | -0.0015041 | -44.8 | 1.4824 | 19.7294 | [5 empty] | -0.015273
-44.7 | -0.0015041 | 1.4918 | -0.033601 | 19.6959 | 1.4918 | 1.7045 | 1.7045 | 1.6883 | 1.4811 | 1e-05 | 1e-05 |  | 1.1827e-07 | -0.0015041 | -44.7 | 1.4811 | 19.6959 | [5 empty] | -0.015273
-44.6 | -0.001467 | 1.4884 | 0 | 19.6624 | 1.4884 | 1.6994 | 1.6994 | 1.6862 | 1.4797 | 1e-05 | 0.018194 |  | 1.2229e-07 | -0.001467 | -44.6 | 1.4797 | 19.6624 | [5 empty] | -0.012219
-44.5 | -0.0015041 | 1.4884 | 0 | 19.6289 | 1.4884 | 1.6994 | 1.6994 | 1.6841 | 1.4783 | 1e-05 | 1e-05 |  | 1.2646e-07 | -0.0015041 | -44.5 | 1.4783 | 19.6289 | [5 empty] | -0.012219
-44.4 | -0.001467 | 1.4884 | -0.033601 | 19.5954 | 1.4884 | 1.6994 | 1.6994 | 1.682 | 1.477 | 1e-05 | 0.015455 |  | 1.3077e-07 | -0.001467 | -44.4 | 1.477 | 19.5954 | [5 empty] | -0.0061093
-44.3 | -0.0015041 | 1.485 | -0.033601 | 19.5619 | 1.485 | 1.6943 | 1.6943 | 1.6799 | 1.4756 | 1e-05 | 1e-05 |  | 1.3522e-07 | -0.0015041 | -44.3 | 1.4756 | 19.5619 | [5 empty] | -0.009164
-44.2 | -0.0015041 | 1.4817 | 0 | 19.5284 | 1.4817 | 1.6892 | 1.6892 | 1.6779 | 1.4742 | 1e-05 | 1e-05 |  | 1.3983e-07 | -0.0015041 | -44.2 | 1.4742 | 19.5284 | [5 empty] | -0.012219
-44.1 | -0.001467 | 1.4817 | 0 | 19.4949 | 1.4817 | 1.6892 | 1.6892 | 1.6758 | 1.4729 | 1e-05 | 0.011344 |  | 1.4459e-07 | -0.001467 | -44.1 | 1.4729 | 19.4949 | [5 empty] | -0.009164
-44 | -0.0015041 | 1.4817 | 0 | 19.4614 | 1.4817 | 1.6892 | 1.6892 | 1.6737 | 1.4715 | 1e-05 | 1e-05 |  | 1.4952e-07 | -0.0015041 | -44 | 1.4715 | 19.4614 | [5 empty] | -0.012219
-43.9 | -0.0015041 | 1.4817 | 0.033601 | 19.4279 | 1.4817 | 1.6892 | 1.6892 | 1.6716 | 1.4701 | 1e-05 | 1e-05 |  | 1.5461e-07 | -0.0015041 | -43.9 | 1.4701 | 19.4279 | [5 empty] | -0.012219
-43.8 | -0.001467 | 1.485 | -0.033601 | 19.3944 | 1.485 | 1.6943 | 1.6943 | 1.6695 | 1.4687 | 1e-05 | 0.0072288 |  | 1.5988e-07 | -0.001467 | -43.8 | 1.4687 | 19.3944 | [5 empty] | -0.009164
-43.7 | -0.001467 | 1.4817 | -0.033601 | 19.3609 | 1.4817 | 1.6892 | 1.6892 | 1.6675 | 1.4674 | 1e-05 | 0.0058563 |  | 1.6532e-07 | -0.001467 | -43.7 | 1.4674 | 19.3609 | [5 empty] | -0.0061093
-43.6 | -0.0015041 | 1.4783 | 0 | 19.3274 | 1.4783 | 1.6841 | 1.6841 | 1.6654 | 1.466 | 1e-05 | 1e-05 |  | 1.7095e-07 | -0.0015041 | -43.6 | 1.466 | 19.3274 | [5 empty] | -0.009164
-43.5 | -0.001467 | 1.4783 | -0.033601 | 19.2939 | 1.4783 | 1.6841 | 1.6841 | 1.6633 | 1.4646 | 1e-05 | 0.0031099 |  | 1.7678e-07 | -0.001467 | -43.5 | 1.4646 | 19.2939 | [5 empty] | -0.009164
-43.4 | -0.0015041 | 1.475 | 0 | 19.2604 | 1.475 | 1.679 | 1.679 | 1.6612 | 1.4633 | 1e-05 | 1e-05 |  | 1.828e-07 | -0.0015041 | -43.4 | 1.4633 | 19.2604 | [5 empty] | -0.01529
-43.3 | -0.0015041 | 1.475 | 0 | 19.2269 | 1.475 | 1.679 | 1.679 | 1.6591 | 1.4619 | 1e-05 | 1e-05 |  | 1.8903e-07 | -0.0015041 | -43.3 | 1.4619 | 19.2269 | [5 empty] | -0.021399
-43.2 | -0.001467 | 1.475 | 0 | 19.1934 | 1.475 | 1.679 | 1.679 | 1.6571 | 1.4605 | 1e-05 | 1e-05 |  | 1.9547e-07 | -0.001467 | -43.2 | 1.4605 | 19.1934 | [5 empty] | -0.018344
-43.1 | -0.001467 | 1.475 | -0.033601 | 19.1599 | 1.475 | 1.679 | 1.679 | 1.655 | 1.4591 | 1e-05 | 1e-05 |  | 2.0212e-07 | -0.001467 | -43.1 | 1.4591 | 19.1599 | [5 empty] | -0.01529
-43 | -0.001467 | 1.4716 | 0 | 19.1265 | 1.4716 | 1.6739 | 1.6739 | 1.6529 | 1.4578 | 1e-05 | 1e-05 |  | 2.0901e-07 | -0.001467 | -43 | 1.4578 | 19.1265 | [5 empty] | -0.01529
-42.9 | -0.001467 | 1.4716 | -0.067384 | 19.093 | 1.4716 | 1.6739 | 1.6739 | 1.6508 | 1.4564 | 1e-05 | 1e-05 |  | 2.1613e-07 | -0.001467 | -42.9 | 1.4564 | 19.093 | [5 empty] | -0.012235
-42.8 | -0.0015041 | 1.4649 | -0.033601 | 19.0595 | 1.4649 | 1.6637 | 1.6637 | 1.6487 | 1.455 | 1e-05 | 1e-05 |  | 2.2349e-07 | -0.0015041 | -42.8 | 1.455 | 19.0595 | [5 empty] | -0.01529
-42.7 | -0.001467 | 1.4615 | 0 | 19.026 | 1.4615 | 1.6586 | 1.6586 | 1.6467 | 1.4536 | 1e-05 | 1e-05 |  | 2.311e-07 | -0.001467 | -42.7 | 1.4536 | 19.026 | [5 empty] | -0.01529
-42.6 | -0.001467 | 1.4615 | 0 | 18.9925 | 1.4615 | 1.6586 | 1.6586 | 1.6446 | 1.4522 | 1e-05 | 1e-05 |  | 2.3898e-07 | -0.001467 | -42.6 | 1.4522 | 18.9925 | [5 empty] | -0.01529
-42.5 | -0.001467 | 1.4615 | 0 | 18.959 | 1.4615 | 1.6586 | 1.6586 | 1.6425 | 1.4509 | 1e-05 | 1e-05 |  | 2.4712e-07 | -0.001467 | -42.5 | 1.4509 | 18.959 | [5 empty] | -0.015306
-42.4 | -0.001467 | 1.4615 | 0 | 18.9255 | 1.4615 | 1.6586 | 1.6586 | 1.6404 | 1.4495 | 1e-05 | 1e-05 |  | 2.5553e-07 | -0.001467 | -42.4 | 1.4495 | 18.9255 | [5 empty] | -0.018377
-42.3 | -0.0014297 | 1.4615 | -0.033601 | 18.892 | 1.4615 | 1.6586 | 1.6586 | 1.6383 | 1.4481 | 1e-05 | 0.023756 |  | 2.6424e-07 | -0.0014297 | -42.3 | 1.4481 | 18.892 | [5 empty] | -0.0091804
-42.2 | -0.001467 | 1.4581 | 0 | 18.8585 | 1.4581 | 1.6535 | 1.6535 | 1.6363 | 1.4467 | 1e-05 | 1e-05 |  | 2.7324e-07 | -0.001467 | -42.2 | 1.4467 | 18.8585 | [5 empty] | -0.0091969
-42.1 | -0.001467 | 1.4581 | 0 | 18.825 | 1.4581 | 1.6535 | 1.6535 | 1.6342 | 1.4454 | 1e-05 | 1e-05 |  | 2.8255e-07 | -0.001467 | -42.1 | 1.4454 | 18.825 | [5 empty] | -0.0091969
-42 | -0.001467 | 1.4581 | -0.033782 | 18.7915 | 1.4581 | 1.6535 | 1.6535 | 1.6321 | 1.444 | 1e-05 | 1e-05 |  | 2.9217e-07 | -0.001467 | -42 | 1.444 | 18.7915 | [5 empty] | -0.012268
-41.9 | -0.001467 | 1.4548 | -0.033782 | 18.758 | 1.4548 | 1.6484 | 1.6484 | 1.63 | 1.4426 | 1e-05 | 1e-05 |  | 3.0212e-07 | -0.001467 | -41.9 | 1.4426 | 18.758 | [5 empty] | -0.012268
-41.8 | -0.001467 | 1.4514 | 0.033782 | 18.7245 | 1.4514 | 1.6433 | 1.6433 | 1.6279 | 1.4412 | 1e-05 | 1e-05 |  | 3.1241e-07 | -0.001467 | -41.8 | 1.4412 | 18.7245 | [5 empty] | -0.012268
-41.7 | -0.001467 | 1.4548 | -0.033782 | 18.691 | 1.4548 | 1.6484 | 1.6484 | 1.6259 | 1.4398 | 1e-05 | 1e-05 |  | 3.2306e-07 | -0.001467 | -41.7 | 1.4398 | 18.691 | [5 empty] | -0.012285
-41.6 | -0.001467 | 1.4514 | 0 | 18.6575 | 1.4514 | 1.6433 | 1.6433 | 1.6238 | 1.4385 | 1e-05 | 1e-05 |  | 3.3406e-07 | -0.001467 | -41.6 | 1.4385 | 18.6575 | [5 empty] | -0.015356
-41.5 | -0.001467 | 1.4514 | -0.033782 | 18.624 | 1.4514 | 1.6433 | 1.6433 | 1.6217 | 1.4371 | 1e-05 | 1e-05 |  | 3.4544e-07 | -0.001467 | -41.5 | 1.4371 | 18.624 | [5 empty] | -0.018427
-41.4 | -0.0014297 | 1.448 | 0 | 18.5905 | 1.448 | 1.6382 | 1.6382 | 1.6196 | 1.4357 | 1e-05 | 0.011329 |  | 3.5721e-07 | -0.0014297 | -41.4 | 1.4357 | 18.5905 | [5 empty] | -0.015356
-41.3 | -0.0014297 | 1.448 | 0 | 18.557 | 1.448 | 1.6382 | 1.6382 | 1.6175 | 1.4343 | 1e-05 | 0.0099463 |  | 3.6938e-07 | -0.0014297 | -41.3 | 1.4343 | 18.557 | [5 empty] | -0.012285
-41.2 | -0.001467 | 1.448 | -0.033782 | 18.5235 | 1.448 | 1.6382 | 1.6382 | 1.6155 | 1.4329 | 1e-05 | 1e-05 |  | 3.8196e-07 | -0.001467 | -41.2 | 1.4329 | 18.5235 | [5 empty] | -0.018427
-41.1 | -0.0014297 | 1.4446 | -0.033782 | 18.49 | 1.4446 | 1.6331 | 1.6331 | 1.6134 | 1.4315 | 1e-05 | 0.007179 |  | 3.9497e-07 | -0.0014297 | -41.1 | 1.4315 | 18.49 | [5 empty] | -0.015356
-41 | -0.001467 | 1.4412 | -0.033782 | 18.4565 | 1.4412 | 1.628 | 1.628 | 1.6113 | 1.4302 | 1e-05 | 1e-05 |  | 4.0842e-07 | -0.001467 | -41 | 1.4302 | 18.4565 | [5 empty] | -0.018427
-40.9 | -0.0014297 | 1.4379 | 0 | 18.423 | 1.4379 | 1.6229 | 1.6229 | 1.6092 | 1.4288 | 1e-05 | 0.0044099 |  | 4.2234e-07 | -0.0014297 | -40.9 | 1.4288 | 18.423 | [5 empty] | -0.015356
-40.8 | -0.001467 | 1.4379 | 0 | 18.3895 | 1.4379 | 1.6229 | 1.6229 | 1.6071 | 1.4274 | 1e-05 | 1e-05 |  | 4.3672e-07 | -0.001467 | -40.8 | 1.4274 | 18.3895 | [5 empty] | -0.018427
-40.7 | -0.001467 | 1.4379 | -0.033782 | 18.356 | 1.4379 | 1.6229 | 1.6229 | 1.6051 | 1.426 | 1e-05 | 1e-05 |  | 4.516e-07 | -0.001467 | -40.7 | 1.426 | 18.356 | [5 empty] | -0.021498
-40.6 | -0.0014297 | 1.4345 | 0 | 18.3225 | 1.4345 | 1.6178 | 1.6178 | 1.603 | 1.4246 | 1e-05 | 0.00025288 |  | 4.6698e-07 | -0.0014297 | -40.6 | 1.4246 | 18.3225 | [5 empty] | -0.018427
-40.5 | -0.0014297 | 1.4345 | -0.033782 | 18.289 | 1.4345 | 1.6178 | 1.6178 | 1.6009 | 1.4232 | 1e-05 | 1e-05 |  | 4.8289e-07 | -0.0014297 | -40.5 | 1.4232 | 18.289 | [5 empty] | -0.015356
-40.4 | -0.0014297 | 1.4311 | 0 | 18.2555 | 1.4311 | 1.6127 | 1.6127 | 1.5988 | 1.4218 | 1e-05 | 1e-05 |  | 4.9934e-07 | -0.0014297 | -40.4 | 1.4218 | 18.2555 | [5 empty] | -0.012285
-40.3 | -0.0014297 | 1.4311 | -0.033782 | 18.2221 | 1.4311 | 1.6127 | 1.6127 | 1.5967 | 1.4205 | 1e-05 | 1e-05 |  | 5.1635e-07 | -0.0014297 | -40.3 | 1.4205 | 18.2221 | [5 empty] | -0.012285
-40.2 | -0.0014297 | 1.4277 | -0.033782 | 18.1886 | 1.4277 | 1.6077 | 1.6077 | 1.5947 | 1.4191 | 1e-05 | 1e-05 |  | 5.3394e-07 | -0.0014297 | -40.2 | 1.4191 | 18.1886 | [5 empty] | -0.015339
-40.1 | -0.0014297 | 1.4244 | 0 | 18.1551 | 1.4244 | 1.6026 | 1.6026 | 1.5926 | 1.4177 | 1e-05 | 1e-05 |  | 5.5212e-07 | -0.0014297 | -40.1 | 1.4177 | 18.1551 | [5 empty] | -0.015323
-40 | -0.0014297 | 1.4244 | 0 | 18.1216 | 1.4244 | 1.6026 | 1.6026 | 1.5905 | 1.4163 | 1e-05 | 1e-05 |  | 5.7093e-07 | -0.0014297 | -40 | 1.4163 | 18.1216 | [5 empty] | -0.018377
-39.9 | -0.0014297 | 1.4244 | 0 | 18.0881 | 1.4244 | 1.6026 | 1.6026 | 1.5884 | 1.4149 | 1e-05 | 1e-05 |  | 5.9038e-07 | -0.0014297 | -39.9 | 1.4149 | 18.0881 | [5 empty] | -0.015306
-39.8 | -0.0014297 | 1.4244 | 0 | 18.0546 | 1.4244 | 1.6026 | 1.6026 | 1.5863 | 1.4135 | 1e-05 | 1e-05 |  | 6.1049e-07 | -0.0014297 | -39.8 | 1.4135 | 18.0546 | [5 empty] | -0.015306
-39.7 | -0.0014297 | 1.4244 | -0.033601 | 18.0211 | 1.4244 | 1.6026 | 1.6026 | 1.5843 | 1.4121 | 1e-05 | 1e-05 |  | 6.3128e-07 | -0.0014297 | -39.7 | 1.4121 | 18.0211 | [5 empty] | -0.012235
-39.6 | -0.0014297 | 1.421 | -0.033601 | 17.9876 | 1.421 | 1.5976 | 1.5976 | 1.5822 | 1.4107 | 1e-05 | 1e-05 |  | 6.5279e-07 | -0.0014297 | -39.6 | 1.4107 | 17.9876 | [5 empty] | -0.012219
-39.5 | -0.0014297 | 1.4176 | -0.033601 | 17.9541 | 1.4176 | 1.5925 | 1.5925 | 1.5801 | 1.4093 | 1e-05 | 1e-05 |  | 6.7502e-07 | -0.0014297 | -39.5 | 1.4093 | 17.9541 | [5 empty] | -0.015273
-39.4 | -0.0014297 | 1.4143 | 0 | 17.9206 | 1.4143 | 1.5875 | 1.5875 | 1.578 | 1.4079 | 1e-05 | 1e-05 |  | 6.9802e-07 | -0.0014297 | -39.4 | 1.4079 | 17.9206 | [5 empty] | -0.015273
-39.3 | -0.0014297 | 1.4143 | 0 | 17.8871 | 1.4143 | 1.5875 | 1.5875 | 1.5759 | 1.4066 | 1e-05 | 1e-05 |  | 7.218e-07 | -0.0014297 | -39.3 | 1.4066 | 17.8871 | [5 empty] | -0.018328
-39.2 | -0.0014297 | 1.4143 | 0 | 17.8536 | 1.4143 | 1.5875 | 1.5875 | 1.5739 | 1.4052 | 1e-05 | 1e-05 |  | 7.4638e-07 | -0.0014297 | -39.2 | 1.4052 | 17.8536 | [5 empty] | -0.021383
-39.1 | -0.0013926 | 1.4143 | -0.033601 | 17.8201 | 1.4143 | 1.5875 | 1.5875 | 1.5718 | 1.4038 | 1e-05 | 0.016368 |  | 7.7181e-07 | -0.0013926 | -39.1 | 1.4038 | 17.8201 | [5 empty] | -0.015273
-39 | -0.0013926 | 1.4109 | -0.033601 | 17.7866 | 1.4109 | 1.5825 | 1.5825 | 1.5697 | 1.4024 | 1e-05 | 0.014975 |  | 7.981e-07 | -0.0013926 | -39 | 1.4024 | 17.7866 | [5 empty] | -0.012219
-38.9 | -0.0013926 | 1.4076 | 0 | 17.7531 | 1.4076 | 1.5774 | 1.5774 | 1.5676 | 1.401 | 1e-05 | 0.013581 |  | 8.2528e-07 | -0.0013926 | -38.9 | 1.401 | 17.7531 | [5 empty] | -0.009164
-38.8 | -0.0014297 | 1.4076 | -0.033601 | 17.7196 | 1.4076 | 1.5774 | 1.5774 | 1.5655 | 1.3996 | 1e-05 | 1e-05 |  | 8.534e-07 | -0.0014297 | -38.8 | 1.3996 | 17.7196 | [5 empty] | -0.012219
-38.7 | -0.0014297 | 1.4042 | -0.033601 | 17.6861 | 1.4042 | 1.5724 | 1.5724 | 1.5635 | 1.3982 | 1e-05 | 1e-05 |  | 8.8247e-07 | -0.0014297 | -38.7 | 1.3982 | 17.6861 | [5 empty] | -0.012219
-38.6 | -0.0014297 | 1.4008 | 0.033601 | 17.6526 | 1.4008 | 1.5674 | 1.5674 | 1.5614 | 1.3968 | 1e-05 | 1e-05 |  | 9.1253e-07 | -0.0014297 | -38.6 | 1.3968 | 17.6526 | [5 empty] | -0.012219
-38.5 | -0.0013926 | 1.4042 | 0 | 17.6191 | 1.4042 | 1.5724 | 1.5724 | 1.5593 | 1.3954 | 1e-05 | 0.0080008 |  | 9.4361e-07 | -0.0013926 | -38.5 | 1.3954 | 17.6191 | [5 empty] | -0.009164
-38.4 | -0.0013926 | 1.4042 | 0 | 17.5856 | 1.4042 | 1.5724 | 1.5724 | 1.5572 | 1.394 | 1e-05 | 0.0066046 |  | 9.7575e-07 | -0.0013926 | -38.4 | 1.394 | 17.5856 | [5 empty] | -0.0061093
-38.3 | -0.0014297 | 1.4042 | -0.033601 | 17.5521 | 1.4042 | 1.5724 | 1.5724 | 1.5551 | 1.3926 | 1e-05 | 1e-05 |  | 1.009e-06 | -0.0014297 | -38.3 | 1.3926 | 17.5521 | [5 empty] | -0.0061093
-38.2 | -0.0013926 | 1.4008 | 0 | 17.5186 | 1.4008 | 1.5674 | 1.5674 | 1.5531 | 1.3912 | 1e-05 | 0.0038109 |  | 1.0434e-06 | -0.0013926 | -38.2 | 1.3912 | 17.5186 | [5 empty] | -0.0030547
-38.1 | -0.0013926 | 1.4008 | 0 | 17.4851 | 1.4008 | 1.5674 | 1.5674 | 1.551 | 1.3898 | 1e-05 | 0.0024133 |  | 1.0789e-06 | -0.0013926 | -38.1 | 1.3898 | 17.4851 | [5 empty] | -2.17e-16
-38 | -0.0014297 | 1.4008 | 0 | 17.4516 | 1.4008 | 1.5674 | 1.5674 | 1.5489 | 1.3884 | 1e-05 | 1e-05 |  | 1.1157e-06 | -0.0014297 | -38 | 1.3884 | 17.4516 | [5 empty] | -0.0061093
-37.9 | -0.0013926 | 1.4008 | 0 | 17.4181 | 1.4008 | 1.5674 | 1.5674 | 1.5468 | 1.387 | 1e-05 | 1e-05 |  | 1.1537e-06 | -0.0013926 | -37.9 | 1.387 | 17.4181 | [5 empty] | -0.0061093
-37.8 | -0.0013926 | 1.4008 | 0 | 17.3846 | 1.4008 | 1.5674 | 1.5674 | 1.5447 | 1.3856 | 1e-05 | 1e-05 |  | 1.193e-06 | -0.0013926 | -37.8 | 1.3856 | 17.3846 | [5 empty] | -0.009164
-37.7 | -0.0013926 | 1.4008 | 0 | 17.3511 | 1.4008 | 1.5674 | 1.5674 | 1.5427 | 1.3842 | 1e-05 | 1e-05 |  | 1.2336e-06 | -0.0013926 | -37.7 | 1.3842 | 17.3511 | [5 empty] | -0.0061093
-37.6 | -0.0014297 | 1.4008 | 0 | 17.3176 | 1.4008 | 1.5674 | 1.5674 | 1.5406 | 1.3828 | 1e-05 | 1e-05 |  | 1.2756e-06 | -0.0014297 | -37.6 | 1.3828 | 17.3176 | [5 empty] | -0.009164
-37.5 | -0.0014297 | 1.4008 | -0.033601 | 17.2842 | 1.4008 | 1.5674 | 1.5674 | 1.5385 | 1.3814 | 1e-05 | 1e-05 |  | 1.3191e-06 | -0.0014297 | -37.5 | 1.3814 | 17.2842 | [5 empty] | -0.012219
-37.4 | -0.0013926 | 1.3975 | 0 | 17.2507 | 1.3975 | 1.5624 | 1.5624 | 1.5364 | 1.38 | 1e-05 | 1e-05 |  | 1.364e-06 | -0.0013926 | -37.4 | 1.38 | 17.2507 | [5 empty] | -0.012219
-37.3 | -0.0013926 | 1.3975 | -0.033601 | 17.2172 | 1.3975 | 1.5624 | 1.5624 | 1.5343 | 1.3786 | 1e-05 | 1e-05 |  | 1.4105e-06 | -0.0013926 | -37.3 | 1.3786 | 17.2172 | [5 empty] | -0.015273
-37.2 | -0.0014297 | 1.3941 | 0 | 17.1837 | 1.3941 | 1.5574 | 1.5574 | 1.5323 | 1.3772 | 1e-05 | 1e-05 |  | 1.4585e-06 | -0.0014297 | -37.2 | 1.3772 | 17.1837 | [5 empty] | -0.021383
-37.1 | -0.0013926 | 1.3941 | -0.033601 | 17.1502 | 1.3941 | 1.5574 | 1.5574 | 1.5302 | 1.3758 | 1e-05 | 1e-05 |  | 1.5082e-06 | -0.0013926 | -37.1 | 1.3758 | 17.1502 | [5 empty] | -0.024437
-37 | -0.0013926 | 1.3908 | -0.033601 | 17.1167 | 1.3908 | 1.5524 | 1.5524 | 1.5281 | 1.3744 | 1e-05 | 1e-05 |  | 1.5596e-06 | -0.0013926 | -37 | 1.3744 | 17.1167 | [5 empty] | -0.024437
-36.9 | -0.0013926 | 1.3874 | 0 | 17.0832 | 1.3874 | 1.5474 | 1.5474 | 1.526 | 1.373 | 1e-05 | 1e-05 |  | 1.6127e-06 | -0.0013926 | -36.9 | 1.373 | 17.0832 | [5 empty] | -0.021383
-36.8 | -0.0013926 | 1.3874 | -0.033601 | 17.0497 | 1.3874 | 1.5474 | 1.5474 | 1.5239 | 1.3716 | 1e-05 | 1e-05 |  | 1.6676e-06 | -0.0013926 | -36.8 | 1.3716 | 17.0497 | [5 empty] | -0.024437
-36.7 | -0.0013555 | 1.384 | -0.067203 | 17.0162 | 1.384 | 1.5424 | 1.5424 | 1.5219 | 1.3702 | 1e-05 | 0.019759 |  | 1.7244e-06 | -0.0013555 | -36.7 | 1.3702 | 17.0162 | [5 empty] | -0.018328
-36.6 | -0.0013926 | 1.3773 | -0.033601 | 16.9827 | 1.3773 | 1.5324 | 1.5324 | 1.5198 | 1.3688 | 1e-05 | 1e-05 |  | 1.7832e-06 | -0.0013926 | -36.6 | 1.3688 | 16.9827 | [5 empty] | -0.021383
-36.5 | -0.0013926 | 1.374 | 0 | 16.9492 | 1.374 | 1.5274 | 1.5274 | 1.5177 | 1.3674 | 1e-05 | 1e-05 |  | 1.8439e-06 | -0.0013926 | -36.5 | 1.3674 | 16.9492 | [5 empty] | -0.021383
-36.4 | -0.0013926 | 1.374 | 0 | 16.9157 | 1.374 | 1.5274 | 1.5274 | 1.5156 | 1.366 | 1e-05 | 1e-05 |  | 1.9067e-06 | -0.0013926 | -36.4 | 1.366 | 16.9157 | [5 empty] | -0.021383
-36.3 | -0.0013926 | 1.374 | -0.033601 | 16.8822 | 1.374 | 1.5274 | 1.5274 | 1.5135 | 1.3646 | 1e-05 | 1e-05 |  | 1.9717e-06 | -0.0013926 | -36.3 | 1.3646 | 16.8822 | [5 empty] | -0.021383
-36.2 | -0.0013555 | 1.3706 | 0.033601 | 16.8487 | 1.3706 | 1.5225 | 1.5225 | 1.5115 | 1.3632 | 1e-05 | 0.012731 |  | 2.0388e-06 | -0.0013555 | -36.2 | 1.3632 | 16.8487 | [5 empty] | -0.018328
-36.1 | -0.0013555 | 1.374 | -0.033601 | 16.8152 | 1.374 | 1.5274 | 1.5274 | 1.5094 | 1.3618 | 1e-05 | 0.011324 |  | 2.1083e-06 | -0.0013555 | -36.1 | 1.3618 | 16.8152 | [5 empty] | -0.012219
-36 | -0.0013555 | 1.3706 | -0.033601 | 16.7817 | 1.3706 | 1.5225 | 1.5225 | 1.5073 | 1.3604 | 1e-05 | 0.0099159 |  | 2.1801e-06 | -0.0013555 | -36 | 1.3604 | 16.7817 | [5 empty] | -0.0061093
-35.9 | -0.0013926 | 1.3672 | -0.033601 | 16.7482 | 1.3672 | 1.5175 | 1.5175 | 1.5052 | 1.3589 | 1e-05 | 1e-05 |  | 2.2544e-06 | -0.0013926 | -35.9 | 1.3589 | 16.7482 | [5 empty] | -0.009164
-35.8 | -0.0013926 | 1.3639 | 0 | 16.7147 | 1.3639 | 1.5125 | 1.5125 | 1.5031 | 1.3575 | 1e-05 | 1e-05 |  | 2.3311e-06 | -0.0013926 | -35.8 | 1.3575 | 16.7147 | [5 empty] | -0.012219
-35.7 | -0.0013555 | 1.3639 | 0 | 16.6812 | 1.3639 | 1.5125 | 1.5125 | 1.5011 | 1.3561 | 1e-05 | 0.0056902 |  | 2.4106e-06 | -0.0013555 | -35.7 | 1.3561 | 16.6812 | [5 empty] | -0.009164
-35.6 | -0.0013926 | 1.3639 | 0 | 16.6477 | 1.3639 | 1.5125 | 1.5125 | 1.499 | 1.3547 | 1e-05 | 1e-05 |  | 2.4927e-06 | -0.0013926 | -35.6 | 1.3547 | 16.6477 | [5 empty] | -0.018344
-35.5 | -0.0013555 | 1.3639 | 0.033601 | 16.6142 | 1.3639 | 1.5125 | 1.5125 | 1.4969 | 1.3533 | 1e-05 | 0.0028707 |  | 2.5776e-06 | -0.0013555 | -35.5 | 1.3533 | 16.6142 | [5 empty] | -0.01529
-35.4 | -0.0013555 | 1.3672 | -0.033601 | 16.5807 | 1.3672 | 1.5175 | 1.5175 | 1.4948 | 1.3519 | 1e-05 | 0.0014601 |  | 2.6654e-06 | -0.0013555 | -35.4 | 1.3519 | 16.5807 | [5 empty] | -0.012235
-35.3 | -0.0013555 | 1.3639 | -0.033601 | 16.5472 | 1.3639 | 1.5125 | 1.5125 | 1.4927 | 1.3505 | 1e-05 | 4.9138e-05 |  | 2.7562e-06 | -0.0013555 | -35.3 | 1.3505 | 16.5472 | [5 empty] | -0.012252
-35.2 | -0.0013926 | 1.3605 | 0 | 16.5137 | 1.3605 | 1.5075 | 1.5075 | 1.4907 | 1.3491 | 1e-05 | 1e-05 |  | 2.8501e-06 | -0.0013926 | -35.2 | 1.3491 | 16.5137 | [5 empty] | -0.018377
-35.1 | -0.0013555 | 1.3605 | -0.067384 | 16.4802 | 1.3605 | 1.5075 | 1.5075 | 1.4886 | 1.3477 | 1e-05 | 1e-05 |  | 2.9471e-06 | -0.0013555 | -35.1 | 1.3477 | 16.4802 | [5 empty] | -0.018377
-35 | -0.0013555 | 1.3538 | 0 | 16.4467 | 1.3538 | 1.4976 | 1.4976 | 1.4865 | 1.3462 | 1e-05 | 1e-05 |  | 3.0475e-06 | -0.0013555 | -35 | 1.3462 | 16.4467 | [5 empty] | -0.018377
-34.9 | -0.0013926 | 1.3538 | 0 | 16.4132 | 1.3538 | 1.4976 | 1.4976 | 1.4844 | 1.3448 | 1e-05 | 1e-05 |  | 3.1513e-06 | -0.0013926 | -34.9 | 1.3448 | 16.4132 | [5 empty] | -0.024487
-34.8 | -0.0013555 | 1.3538 | -0.033782 | 16.3797 | 1.3538 | 1.4976 | 1.4976 | 1.4823 | 1.3434 | 1e-05 | 1e-05 |  | 3.2587e-06 | -0.0013555 | -34.8 | 1.3434 | 16.3797 | [5 empty] | -0.021432
-34.7 | -0.0013555 | 1.3504 | -0.067384 | 16.3463 | 1.3504 | 1.4926 | 1.4926 | 1.4803 | 1.342 | 1e-05 | 1e-05 |  | 3.3697e-06 | -0.0013555 | -34.7 | 1.342 | 16.3463 | [5 empty] | -0.018377
-34.6 | -0.0013555 | 1.3437 | 0 | 16.3128 | 1.3437 | 1.4827 | 1.4827 | 1.4782 | 1.3406 | 1e-05 | 1e-05 |  | 3.4845e-06 | -0.0013555 | -34.6 | 1.3406 | 16.3128 | [5 empty] | -0.018377
-34.5 | -0.0013182 | 1.3437 | 0 | 16.2793 | 1.3437 | 1.4827 | 1.4827 | 1.4761 | 1.3392 | 1e-05 | 0.025904 |  | 3.6032e-06 | -0.0013182 | -34.5 | 1.3392 | 16.2793 | [5 empty] | -0.012252
-34.4 | -0.0013555 | 1.3437 | -0.033601 | 16.2458 | 1.3437 | 1.4827 | 1.4827 | 1.474 | 1.3378 | 1e-05 | 1e-05 |  | 3.7259e-06 | -0.0013555 | -34.4 | 1.3378 | 16.2458 | [5 empty] | -0.012252
-34.3 | -0.0013555 | 1.3403 | 0 | 16.2123 | 1.3403 | 1.4778 | 1.4778 | 1.4719 | 1.3363 | 1e-05 | 1e-05 |  | 3.8528e-06 | -0.0013555 | -34.3 | 1.3363 | 16.2123 | [5 empty] | -0.012252
-34.2 | -0.0013182 | 1.3403 | 0 | 16.1788 | 1.3403 | 1.4778 | 1.4778 | 1.4699 | 1.3349 | 1e-05 | 0.021656 |  | 3.9841e-06 | -0.0013182 | -34.2 | 1.3349 | 16.1788 | [5 empty] | -0.0091804
-34.1 | -0.0013182 | 1.3403 | 0 | 16.1453 | 1.3403 | 1.4778 | 1.4778 | 1.4678 | 1.3335 | 1e-05 | 0.020239 |  | 4.1198e-06 | -0.0013182 | -34.1 | 1.3335 | 16.1453 | [5 empty] | 1.6467e-05
-34 | -0.0013555 | 1.3403 | 0 | 16.1118 | 1.3403 | 1.4778 | 1.4778 | 1.4657 | 1.3321 | 1e-05 | 1e-05 |  | 4.2601e-06 | -0.0013555 | -34 | 1.3321 | 16.1118 | [5 empty] | 1.6467e-05
-33.9 | -0.0013555 | 1.3403 | 0 | 16.0783 | 1.3403 | 1.4778 | 1.4778 | 1.4636 | 1.3307 | 1e-05 | 1e-05 |  | 4.4052e-06 | -0.0013555 | -33.9 | 1.3307 | 16.0783 | [5 empty] | -0.0030547
-33.8 | -0.0013555 | 1.3403 | 0 | 16.0448 | 1.3403 | 1.4778 | 1.4778 | 1.4615 | 1.3293 | 1e-05 | 1e-05 |  | 4.5553e-06 | -0.0013555 | -33.8 | 1.3293 | 16.0448 | [5 empty] | 0
-33.7 | -0.0013555 | 1.3403 | 0 | 16.0113 | 1.3403 | 1.4778 | 1.4778 | 1.4595 | 1.3278 | 1e-05 | 1e-05 |  | 4.7105e-06 | -0.0013555 | -33.7 | 1.3278 | 16.0113 | [5 empty] | -0.0030711
-33.6 | -0.0013555 | 1.3403 | 0.033782 | 15.9778 | 1.3403 | 1.4778 | 1.4778 | 1.4574 | 1.3264 | 1e-05 | 1e-05 |  | 4.8709e-06 | -0.0013555 | -33.6 | 1.3264 | 15.9778 | [5 empty] | -0.0061423
-33.5 | -0.0013555 | 1.3437 | 0 | 15.9443 | 1.3437 | 1.4827 | 1.4827 | 1.4553 | 1.325 | 1e-05 | 1e-05 |  | 5.0368e-06 | -0.0013555 | -33.5 | 1.325 | 15.9443 | [5 empty] | -0.0092134
-33.4 | -0.0013182 | 1.3437 | -0.033782 | 15.9108 | 1.3437 | 1.4827 | 1.4827 | 1.4532 | 1.3236 | 1e-05 | 0.010306 |  | 5.2084e-06 | -0.0013182 | -33.4 | 1.3236 | 15.9108 | [5 empty] | -0.0092134
-33.3 | -0.0013555 | 1.3403 | 0 | 15.8773 | 1.3403 | 1.4778 | 1.4778 | 1.4511 | 1.3222 | 1e-05 | 1e-05 |  | 5.3858e-06 | -0.0013555 | -33.3 | 1.3222 | 15.8773 | [5 empty] | -0.012285
-33.2 | -0.0013555 | 1.3403 | -0.033782 | 15.8438 | 1.3403 | 1.4778 | 1.4778 | 1.4491 | 1.3207 | 1e-05 | 1e-05 |  | 5.5693e-06 | -0.0013555 | -33.2 | 1.3207 | 15.8438 | [5 empty] | -0.015356
-33.1 | -0.0013182 | 1.3369 | -0.033782 | 15.8103 | 1.3369 | 1.4728 | 1.4728 | 1.447 | 1.3193 | 1e-05 | 0.0060416 |  | 5.759e-06 | -0.0013182 | -33.1 | 1.3193 | 15.8103 | [5 empty] | -0.015356
-33 | -0.0013555 | 1.3335 | -0.033782 | 15.7768 | 1.3335 | 1.4678 | 1.4678 | 1.4449 | 1.3179 | 1e-05 | 1e-05 |  | 5.9552e-06 | -0.0013555 | -33 | 1.3179 | 15.7768 | [5 empty] | -0.021498
-32.9 | -0.0013555 | 1.3302 | 0 | 15.7433 | 1.3302 | 1.4629 | 1.4629 | 1.4428 | 1.3165 | 1e-05 | 1e-05 |  | 6.158e-06 | -0.0013555 | -32.9 | 1.3165 | 15.7433 | [5 empty] | -0.024569
-32.8 | -0.0013182 | 1.3302 | -0.033782 | 15.7098 | 1.3302 | 1.4629 | 1.4629 | 1.4408 | 1.315 | 1e-05 | 0.0017725 |  | 6.3678e-06 | -0.0013182 | -32.8 | 1.315 | 15.7098 | [5 empty] | -0.024553
-32.7 | -0.0013555 | 1.3268 | -0.033782 | 15.6763 | 1.3268 | 1.4579 | 1.4579 | 1.4387 | 1.3136 | 1e-05 | 1e-05 |  | 6.5847e-06 | -0.0013555 | -32.7 | 1.3136 | 15.6763 | [5 empty] | -0.027624
-32.6 | -0.0013182 | 1.3234 | 0 | 15.6428 | 1.3234 | 1.453 | 1.453 | 1.4366 | 1.3122 | 1e-05 | 1e-05 |  | 6.809e-06 | -0.0013182 | -32.6 | 1.3122 | 15.6428 | [5 empty] | -0.024553
-32.5 | -0.0013182 | 1.3234 | -0.033782 | 15.6093 | 1.3234 | 1.453 | 1.453 | 1.4345 | 1.3108 | 1e-05 | 1e-05 |  | 7.041e-06 | -0.0013182 | -32.5 | 1.3108 | 15.6093 | [5 empty] | -0.024536
-32.4 | -0.0013182 | 1.32 | -0.033782 | 15.5758 | 1.32 | 1.448 | 1.448 | 1.4324 | 1.3093 | 1e-05 | 1e-05 |  | 7.2808e-06 | -0.0013182 | -32.4 | 1.3093 | 15.5758 | [5 empty] | -0.02452
-32.3 | -0.0013182 | 1.3167 | -0.033601 | 15.5423 | 1.3167 | 1.4431 | 1.4431 | 1.4304 | 1.3079 | 1e-05 | 1e-05 |  | 7.5288e-06 | -0.0013182 | -32.3 | 1.3079 | 15.5423 | [5 empty] | -0.02452
-32.2 | -0.0013182 | 1.3133 | -0.033782 | 15.5088 | 1.3133 | 1.4382 | 1.4382 | 1.4283 | 1.3065 | 1e-05 | 1e-05 |  | 7.7853e-06 | -0.0013182 | -32.2 | 1.3065 | 15.5088 | [5 empty] | -0.021448
-32.1 | -0.0013182 | 1.3099 | 0 | 15.4753 | 1.3099 | 1.4333 | 1.4333 | 1.4262 | 1.3051 | 1e-05 | 1e-05 |  | 8.0505e-06 | -0.0013182 | -32.1 | 1.3051 | 15.4753 | [5 empty] | -0.021432
-32 | -0.0013182 | 1.3099 | -0.033601 | 15.4418 | 1.3099 | 1.4333 | 1.4333 | 1.4241 | 1.3036 | 1e-05 | 1e-05 |  | 8.3247e-06 | -0.0013182 | -32 | 1.3036 | 15.4418 | [5 empty] | -0.024487
-31.9 | -0.0013182 | 1.3066 | -0.033601 | 15.4084 | 1.3066 | 1.4284 | 1.4284 | 1.422 | 1.3022 | 1e-05 | 1e-05 |  | 8.6083e-06 | -0.0013182 | -31.9 | 1.3022 | 15.4084 | [5 empty] | -0.021416
-31.8 | -0.0013182 | 1.3032 | 0 | 15.3749 | 1.3032 | 1.4235 | 1.4235 | 1.42 | 1.3008 | 1e-05 | 1e-05 |  | 8.9015e-06 | -0.0013182 | -31.8 | 1.3008 | 15.3749 | [5 empty] | -0.021399
-31.7 | -0.0012811 | 1.3032 | 0 | 15.3414 | 1.3032 | 1.4235 | 1.4235 | 1.4179 | 1.2994 | 1e-05 | 0.023041 |  | 9.2047e-06 | -0.0012811 | -31.7 | 1.2994 | 15.3414 | [5 empty] | -0.01529
-31.6 | -0.0013182 | 1.3032 | -0.033601 | 15.3079 | 1.3032 | 1.4235 | 1.4235 | 1.4158 | 1.2979 | 1e-05 | 1e-05 |  | 9.5183e-06 | -0.0013182 | -31.6 | 1.2979 | 15.3079 | [5 empty] | -0.015273
-31.5 | -0.0013182 | 1.2998 | -0.033601 | 15.2744 | 1.2998 | 1.4186 | 1.4186 | 1.4137 | 1.2965 | 1e-05 | 1e-05 |  | 9.8425e-06 | -0.0013182 | -31.5 | 1.2965 | 15.2744 | [5 empty] | -0.015273
-31.4 | -0.0012811 | 1.2965 | 0 | 15.2409 | 1.2965 | 1.4137 | 1.4137 | 1.4116 | 1.2951 | 1e-05 | 0.01875 |  | 1.0178e-05 | -0.0012811 | -31.4 | 1.2951 | 15.2409 | [5 empty] | -0.012219
-31.3 | -0.0012811 | 1.2965 | -0.033601 | 15.2074 | 1.2965 | 1.4137 | 1.4137 | 1.4096 | 1.2936 | 1e-05 | 0.017318 |  | 1.0524e-05 | -0.0012811 | -31.3 | 1.2936 | 15.2074 | [5 empty] | -0.0061093
-31.2 | -0.0013182 | 1.2931 | 0.033601 | 15.1739 | 1.2931 | 1.4088 | 1.4088 | 1.4075 | 1.2922 | 1e-05 | 1e-05 |  | 1.0883e-05 | -0.0013182 | -31.2 | 1.2922 | 15.1739 | [5 empty] | -0.009164
-31.1 | -0.0013182 | 1.2965 | -0.033601 | 15.1404 | 1.2965 | 1.4137 | 1.4137 | 1.4054 | 1.2908 | 1e-05 | 1e-05 |  | 1.1254e-05 | -0.0013182 | -31.1 | 1.2908 | 15.1404 | [5 empty] | -0.009164
-31 | -0.0012811 | 1.2931 | 0 | 15.1069 | 1.2931 | 1.4088 | 1.4088 | 1.4033 | 1.2893 | 1e-05 | 0.013021 |  | 1.1637e-05 | -0.0012811 | -31 | 1.2893 | 15.1069 | [5 empty] | -0.0030547
-30.9 | -0.0012811 | 1.2931 | 0 | 15.0734 | 1.2931 | 1.4088 | 1.4088 | 1.4012 | 1.2879 | 1e-05 | 0.011587 |  | 1.2033e-05 | -0.0012811 | -30.9 | 1.2879 | 15.0734 | [5 empty] | 0
-30.8 | -0.0013182 | 1.2931 | 0.033601 | 15.0399 | 1.2931 | 1.4088 | 1.4088 | 1.3992 | 1.2865 | 1e-05 | 1e-05 |  | 1.2443e-05 | -0.0013182 | -30.8 | 1.2865 | 15.0399 | [5 empty] | -0.0030547
-30.7 | -0.0012811 | 1.2965 | -0.033601 | 15.0064 | 1.2965 | 1.4137 | 1.4137 | 1.3971 | 1.285 | 1e-05 | 0.0087189 |  | 1.2867e-05 | -0.0012811 | -30.7 | 1.285 | 15.0064 | [5 empty] | 0.0030547
-30.6 | -0.0013182 | 1.2931 | 0 | 14.9729 | 1.2931 | 1.4088 | 1.4088 | 1.395 | 1.2836 | 1e-05 | 1e-05 |  | 1.3305e-05 | -0.0013182 | -30.6 | 1.2836 | 14.9729 | [5 empty] | 0
-30.5 | -0.0012811 | 1.2931 | 0.033601 | 14.9394 | 1.2931 | 1.4088 | 1.4088 | 1.3929 | 1.2822 | 1e-05 | 0.0058482 |  | 1.3759e-05 | -0.0012811 | -30.5 | 1.2822 | 14.9394 | [5 empty] | 0.0030547
-30.4 | -0.0013182 | 1.2965 | 0 | 14.9059 | 1.2965 | 1.4137 | 1.4137 | 1.3908 | 1.2807 | 1e-05 | 1e-05 |  | 1.4227e-05 | -0.0013182 | -30.4 | 1.2807 | 14.9059 | [5 empty] | -1.085e-16
-30.3 | -0.0012811 | 1.2965 | -0.033601 | 14.8724 | 1.2965 | 1.4137 | 1.4137 | 1.3888 | 1.2793 | 1e-05 | 0.0029753 |  | 1.4712e-05 | -0.0012811 | -30.3 | 1.2793 | 14.8724 | [5 empty] | -1.085e-16
-30.2 | -0.0013182 | 1.2931 | 0.033601 | 14.8389 | 1.2931 | 1.4088 | 1.4088 | 1.3867 | 1.2779 | 1e-05 | 1e-05 |  | 1.5213e-05 | -0.0013182 | -30.2 | 1.2779 | 14.8389 | [5 empty] | -0.0030547
-30.1 | -0.0012811 | 1.2965 | 0 | 14.8054 | 1.2965 | 1.4137 | 1.4137 | 1.3846 | 1.2764 | 1e-05 | 0.00010032 |  | 1.5731e-05 | -0.0012811 | -30.1 | 1.2764 | 14.8054 | [5 empty] | 0
-30 | -0.0013182 | 1.2965 | 0 | 14.7719 | 1.2965 | 1.4137 | 1.4137 | 1.3825 | 1.275 | 1e-05 | 1e-05 |  | 1.6267e-05 | -0.0013182 | -30 | 1.275 | 14.7719 | [5 empty] | -0.0061093
-29.9 | -0.0013182 | 1.2965 | -0.033601 | 14.7384 | 1.2965 | 1.4137 | 1.4137 | 1.3804 | 1.2735 | 1e-05 | 1e-05 |  | 1.6821e-05 | -0.0013182 | -29.9 | 1.2735 | 14.7384 | [5 empty] | -0.012219
-29.8 | -0.0012811 | 1.2931 | 0 | 14.7049 | 1.2931 | 1.4088 | 1.4088 | 1.3784 | 1.2721 | 1e-05 | 1e-05 |  | 1.7394e-05 | -0.0012811 | -29.8 | 1.2721 | 14.7049 | [5 empty] | -0.012219
-29.7 | -0.0012811 | 1.2931 | 0 | 14.6714 | 1.2931 | 1.4088 | 1.4088 | 1.3763 | 1.2707 | 1e-05 | 1e-05 |  | 1.7987e-05 | -0.0012811 | -29.7 | 1.2707 | 14.6714 | [5 empty] | -0.009164
-29.6 | -0.0013182 | 1.2931 | 0 | 14.6379 | 1.2931 | 1.4088 | 1.4088 | 1.3742 | 1.2692 | 1e-05 | 1e-05 |  | 1.86e-05 | -0.0013182 | -29.6 | 1.2692 | 14.6379 | [5 empty] | -0.015273
-29.5 | -0.0013182 | 1.2931 | -0.067203 | 14.6044 | 1.2931 | 1.4088 | 1.4088 | 1.3721 | 1.2678 | 1e-05 | 1e-05 |  | 1.9233e-05 | -0.0013182 | -29.5 | 1.2678 | 14.6044 | [5 empty] | -0.018328
-29.4 | -0.0012811 | 1.2864 | -0.033601 | 14.5709 | 1.2864 | 1.399 | 1.399 | 1.37 | 1.2663 | 1e-05 | 1e-05 |  | 1.9888e-05 | -0.0012811 | -29.4 | 1.2663 | 14.5709 | [5 empty] | -0.018328
-29.3 | -0.0012811 | 1.283 | 0 | 14.5374 | 1.283 | 1.3942 | 1.3942 | 1.368 | 1.2649 | 1e-05 | 1e-05 |  | 2.0566e-05 | -0.0012811 | -29.3 | 1.2649 | 14.5374 | [5 empty] | -0.018328
-29.2 | -0.0012811 | 1.283 | 0 | 14.5039 | 1.283 | 1.3942 | 1.3942 | 1.3659 | 1.2634 | 1e-05 | 1e-05 |  | 2.1266e-05 | -0.0012811 | -29.2 | 1.2634 | 14.5039 | [5 empty] | -0.018328
-29.1 | -0.0013182 | 1.283 | -0.033601 | 14.4705 | 1.283 | 1.3942 | 1.3942 | 1.3638 | 1.262 | 1e-05 | 1e-05 |  | 2.1991e-05 | -0.0013182 | -29.1 | 1.262 | 14.4705 | [5 empty] | -0.021383
-29 | -0.0012811 | 1.2797 | -0.033601 | 14.437 | 1.2797 | 1.3893 | 1.3893 | 1.3617 | 1.2606 | 1e-05 | 1e-05 |  | 2.274e-05 | -0.0012811 | -29 | 1.2606 | 14.437 | [5 empty] | -0.021383
-28.9 | -0.001244 | 1.2763 | 0 | 14.4035 | 1.2763 | 1.3844 | 1.3844 | 1.3596 | 1.2591 | 1e-05 | 0.019766 |  | 2.3514e-05 | -0.001244 | -28.9 | 1.2591 | 14.4035 | [5 empty] | -0.015273
-28.8 | -0.0012811 | 1.2763 | -0.033601 | 14.37 | 1.2763 | 1.3844 | 1.3844 | 1.3576 | 1.2577 | 1e-05 | 1e-05 |  | 2.4315e-05 | -0.0012811 | -28.8 | 1.2577 | 14.37 | [5 empty] | -0.012219
-28.7 | -0.0012811 | 1.273 | 0 | 14.3365 | 1.273 | 1.3796 | 1.3796 | 1.3555 | 1.2562 | 1e-05 | 1e-05 |  | 2.5144e-05 | -0.0012811 | -28.7 | 1.2562 | 14.3365 | [5 empty] | -0.015273
-28.6 | -0.0012811 | 1.273 | -0.033601 | 14.303 | 1.273 | 1.3796 | 1.3796 | 1.3534 | 1.2548 | 1e-05 | 1e-05 |  | 2.6e-05 | -0.0012811 | -28.6 | 1.2548 | 14.303 | [5 empty] | -0.018328
-28.5 | -0.0012811 | 1.2696 | 0 | 14.2695 | 1.2696 | 1.3747 | 1.3747 | 1.3513 | 1.2533 | 1e-05 | 1e-05 |  | 2.6886e-05 | -0.0012811 | -28.5 | 1.2533 | 14.2695 | [5 empty] | -0.018328
-28.4 | -0.0012811 | 1.2696 | 0 | 14.236 | 1.2696 | 1.3747 | 1.3747 | 1.3492 | 1.2519 | 1e-05 | 1e-05 |  | 2.7802e-05 | -0.0012811 | -28.4 | 1.2519 | 14.236 | [5 empty] | -0.018328
-28.3 | -0.0012811 | 1.2696 | 0 | 14.2025 | 1.2696 | 1.3747 | 1.3747 | 1.3472 | 1.2504 | 1e-05 | 1e-05 |  | 2.8749e-05 | -0.0012811 | -28.3 | 1.2504 | 14.2025 | [5 empty] | -0.021383
-28.2 | -0.001244 | 1.2696 | -0.033601 | 14.169 | 1.2696 | 1.3747 | 1.3747 | 1.3451 | 1.249 | 1e-05 | 0.0096402 |  | 2.9728e-05 | -0.001244 | -28.2 | 1.249 | 14.169 | [5 empty] | -0.018328
-28.1 | -0.0012811 | 1.2662 | -0.033601 | 14.1355 | 1.2662 | 1.3699 | 1.3699 | 1.343 | 1.2475 | 1e-05 | 1e-05 |  | 3.0741e-05 | -0.0012811 | -28.1 | 1.2475 | 14.1355 | [5 empty] | -0.021383
-28 | -0.0012811 | 1.2629 | -0.033601 | 14.102 | 1.2629 | 1.365 | 1.365 | 1.3409 | 1.2461 | 1e-05 | 1e-05 |  | 3.1788e-05 | -0.0012811 | -28 | 1.2461 | 14.102 | [5 empty] | -0.021383
-27.9 | -0.0012811 | 1.2595 | -0.033601 | 14.0685 | 1.2595 | 1.3602 | 1.3602 | 1.3388 | 1.2446 | 1e-05 | 1e-05 |  | 3.2871e-05 | -0.0012811 | -27.9 | 1.2446 | 14.0685 | [5 empty] | -0.024437
-27.8 | -0.001244 | 1.2562 | -0.033601 | 14.035 | 1.2562 | 1.3554 | 1.3554 | 1.3368 | 1.2432 | 1e-05 | 0.0038418 |  | 3.399e-05 | -0.001244 | -27.8 | 1.2432 | 14.035 | [5 empty] | -0.024437
-27.7 | -0.0012811 | 1.2528 | 0 | 14.0015 | 1.2528 | 1.3505 | 1.3505 | 1.3347 | 1.2417 | 1e-05 | 1e-05 |  | 3.5148e-05 | -0.0012811 | -27.7 | 1.2417 | 14.0015 | [5 empty] | -0.027492
-27.6 | -0.001244 | 1.2528 | -0.033601 | 13.968 | 1.2528 | 1.3505 | 1.3505 | 1.3326 | 1.2403 | 1e-05 | 0.00093919 |  | 3.6345e-05 | -0.001244 | -27.6 | 1.2403 | 13.968 | [5 empty] | -0.024437
-27.5 | -0.001244 | 1.2494 | -0.033601 | 13.9345 | 1.2494 | 1.3457 | 1.3457 | 1.3305 | 1.2388 | 1e-05 | 1e-05 |  | 3.7583e-05 | -0.001244 | -27.5 | 1.2388 | 13.9345 | [5 empty] | -0.021383
-27.4 | -0.001244 | 1.2461 | -0.033601 | 13.901 | 1.2461 | 1.3409 | 1.3409 | 1.3284 | 1.2374 | 1e-05 | 1e-05 |  | 3.8864e-05 | -0.001244 | -27.4 | 1.2374 | 13.901 | [5 empty] | -0.018328
-27.3 | -0.001244 | 1.2427 | 0 | 13.8675 | 1.2427 | 1.3361 | 1.3361 | 1.3264 | 1.2359 | 1e-05 | 1e-05 |  | 4.0188e-05 | -0.001244 | -27.3 | 1.2359 | 13.8675 | [5 empty] | -0.018344
-27.2 | -0.001244 | 1.2427 | -0.033601 | 13.834 | 1.2427 | 1.3361 | 1.3361 | 1.3243 | 1.2345 | 1e-05 | 1e-05 |  | 4.1556e-05 | -0.001244 | -27.2 | 1.2345 | 13.834 | [5 empty] | -0.018361
-27.1 | -0.001244 | 1.2394 | 0 | 13.8005 | 1.2394 | 1.3313 | 1.3313 | 1.3222 | 1.233 | 1e-05 | 1e-05 |  | 4.2972e-05 | -0.001244 | -27.1 | 1.233 | 13.8005 | [5 empty] | -0.018361
-27 | -0.001244 | 1.2394 | 0 | 13.767 | 1.2394 | 1.3313 | 1.3313 | 1.3201 | 1.2316 | 1e-05 | 1e-05 |  | 4.4436e-05 | -0.001244 | -27 | 1.2316 | 13.767 | [5 empty] | -0.015306
-26.9 | -0.001244 | 1.2394 | 0 | 13.7335 | 1.2394 | 1.3313 | 1.3313 | 1.318 | 1.2301 | 1e-05 | 1e-05 |  | 4.5949e-05 | -0.001244 | -26.9 | 1.2301 | 13.7335 | [5 empty] | -0.012252
-26.8 | -0.001244 | 1.2394 | -0.033782 | 13.7 | 1.2394 | 1.3313 | 1.3313 | 1.316 | 1.2287 | 1e-05 | 1e-05 |  | 4.7515e-05 | -0.001244 | -26.8 | 1.2287 | 13.7 | [5 empty] | -0.0091969
-26.7 | -0.001244 | 1.236 | -0.033782 | 13.6665 | 1.236 | 1.3264 | 1.3264 | 1.3139 | 1.2272 | 1e-05 | 1e-05 |  | 4.9133e-05 | -0.001244 | -26.7 | 1.2272 | 13.6665 | [5 empty] | -0.012268
-26.6 | -0.001244 | 1.2326 | 0 | 13.633 | 1.2326 | 1.3216 | 1.3216 | 1.3118 | 1.2257 | 1e-05 | 1e-05 |  | 5.0807e-05 | -0.001244 | -26.6 | 1.2257 | 13.633 | [5 empty] | -0.012285
-26.5 | -0.001244 | 1.2326 | 0 | 13.5995 | 1.2326 | 1.3216 | 1.3216 | 1.3097 | 1.2243 | 1e-05 | 1e-05 |  | 5.2538e-05 | -0.001244 | -26.5 | 1.2243 | 13.5995 | [5 empty] | -0.015356
-26.4 | -0.001244 | 1.2326 | 0 | 13.566 | 1.2326 | 1.3216 | 1.3216 | 1.3076 | 1.2228 | 1e-05 | 1e-05 |  | 5.4327e-05 | -0.001244 | -26.4 | 1.2228 | 13.566 | [5 empty] | -0.018427
-26.3 | -0.001244 | 1.2326 | 0 | 13.5326 | 1.2326 | 1.3216 | 1.3216 | 1.3056 | 1.2214 | 1e-05 | 1e-05 |  | 5.6178e-05 | -0.001244 | -26.3 | 1.2214 | 13.5326 | [5 empty] | -0.018427
-26.2 | -0.0012067 | 1.2326 | -0.033782 | 13.4991 | 1.2326 | 1.3216 | 1.3216 | 1.3035 | 1.2199 | 1e-05 | 0.017718 |  | 5.8091e-05 | -0.0012067 | -26.2 | 1.2199 | 13.4991 | [5 empty] | -0.015356
-26.1 | -0.0012067 | 1.2292 | -0.033782 | 13.4656 | 1.2292 | 1.3168 | 1.3168 | 1.3014 | 1.2184 | 1e-05 | 0.016257 |  | 6.007e-05 | -0.0012067 | -26.1 | 1.2184 | 13.4656 | [5 empty] | -0.012285
-26 | -0.001244 | 1.2258 | -0.033782 | 13.4321 | 1.2258 | 1.3119 | 1.3119 | 1.2993 | 1.217 | 1e-05 | 1e-05 |  | 6.2116e-05 | -0.001244 | -26 | 1.217 | 13.4321 | [5 empty] | -0.015356
-25.9 | -0.001244 | 1.2225 | -0.033782 | 13.3986 | 1.2225 | 1.3071 | 1.3071 | 1.2972 | 1.2155 | 1e-05 | 1e-05 |  | 6.4232e-05 | -0.001244 | -25.9 | 1.2155 | 13.3986 | [5 empty] | -0.018427
-25.8 | -0.001244 | 1.2191 | 0 | 13.3651 | 1.2191 | 1.3023 | 1.3023 | 1.2952 | 1.2141 | 1e-05 | 1e-05 |  | 6.642e-05 | -0.001244 | -25.8 | 1.2141 | 13.3651 | [5 empty] | -0.021498
-25.7 | -0.001244 | 1.2191 | 0 | 13.3316 | 1.2191 | 1.3023 | 1.3023 | 1.2931 | 1.2126 | 1e-05 | 1e-05 |  | 6.8683e-05 | -0.001244 | -25.7 | 1.2126 | 13.3316 | [5 empty] | -0.024569
-25.6 | -0.0012067 | 1.2191 | 0 | 13.2981 | 1.2191 | 1.3023 | 1.3023 | 1.291 | 1.2111 | 1e-05 | 0.0089476 |  | 7.1022e-05 | -0.0012067 | -25.6 | 1.2111 | 13.2981 | [5 empty] | -0.021498
-25.5 | -0.0012067 | 1.2191 | -0.033782 | 13.2646 | 1.2191 | 1.3023 | 1.3023 | 1.2889 | 1.2097 | 1e-05 | 0.0074839 |  | 7.3442e-05 | -0.0012067 | -25.5 | 1.2097 | 13.2646 | [5 empty] | -0.018427
-25.4 | -0.0012067 | 1.2157 | -0.033782 | 13.2311 | 1.2157 | 1.2975 | 1.2975 | 1.2868 | 1.2082 | 1e-05 | 0.0060196 |  | 7.5943e-05 | -0.0012067 | -25.4 | 1.2082 | 13.2311 | [5 empty] | -0.015356
-25.3 | -0.0012067 | 1.2123 | -0.033782 | 13.1976 | 1.2123 | 1.2927 | 1.2927 | 1.2848 | 1.2067 | 1e-05 | 0.0045547 |  | 7.853e-05 | -0.0012067 | -25.3 | 1.2067 | 13.1976 | [5 empty] | -0.012285
-25.2 | -0.001244 | 1.2089 | -0.033782 | 13.1641 | 1.2089 | 1.2879 | 1.2879 | 1.2827 | 1.2053 | 1e-05 | 1e-05 |  | 8.1205e-05 | -0.001244 | -25.2 | 1.2053 | 13.1641 | [5 empty] | -0.015356
-25.1 | -0.0012067 | 1.2056 | 0 | 13.1306 | 1.2056 | 1.2831 | 1.2831 | 1.2806 | 1.2038 | 1e-05 | 0.0016231 |  | 8.3972e-05 | -0.0012067 | -25.1 | 1.2038 | 13.1306 | [5 empty] | -0.01841
-25 | -0.0012067 | 1.2056 | 0 | 13.0971 | 1.2056 | 1.2831 | 1.2831 | 1.2785 | 1.2023 | 1e-05 | 0.00015641 |  | 8.6832e-05 | -0.0012067 | -25 | 1.2023 | 13.0971 | [5 empty] | -0.021465
-24.9 | -0.0012067 | 1.2056 | 0 | 13.0636 | 1.2056 | 1.2831 | 1.2831 | 1.2764 | 1.2009 | 1e-05 | 1e-05 |  | 8.979e-05 | -0.0012067 | -24.9 | 1.2009 | 13.0636 | [5 empty] | -0.018394
-24.8 | -0.0012067 | 1.2056 | 0 | 13.0301 | 1.2056 | 1.2831 | 1.2831 | 1.2744 | 1.1994 | 1e-05 | 1e-05 |  | 9.2848e-05 | -0.0012067 | -24.8 | 1.1994 | 13.0301 | [5 empty] | -0.018377
-24.7 | -0.0012067 | 1.2056 | -0.033782 | 12.9966 | 1.2056 | 1.2831 | 1.2831 | 1.2723 | 1.1979 | 1e-05 | 1e-05 |  | 9.6011e-05 | -0.0012067 | -24.7 | 1.1979 | 12.9966 | [5 empty] | -0.015306
-24.6 | -0.0012067 | 1.2022 | -0.033601 | 12.9631 | 1.2022 | 1.2783 | 1.2783 | 1.2702 | 1.1965 | 1e-05 | 1e-05 |  | 9.9282e-05 | -0.0012067 | -24.6 | 1.1965 | 12.9631 | [5 empty] | -0.01529
-24.5 | -0.0012067 | 1.1988 | -0.033601 | 12.9296 | 1.1988 | 1.2735 | 1.2735 | 1.2681 | 1.195 | 1e-05 | 1e-05 |  | 0.00010266 | -0.0012067 | -24.5 | 1.195 | 12.9296 | [5 empty] | -0.018344
-24.4 | -0.0012067 | 1.1955 | 0 | 12.8961 | 1.1955 | 1.2688 | 1.2688 | 1.266 | 1.1935 | 1e-05 | 1e-05 |  | 0.00010616 | -0.0012067 | -24.4 | 1.1935 | 12.8961 | [5 empty] | -0.021399
-24.3 | -0.0012067 | 1.1955 | -0.033601 | 12.8626 | 1.1955 | 1.2688 | 1.2688 | 1.264 | 1.1921 | 1e-05 | 1e-05 |  | 0.00010978 | -0.0012067 | -24.3 | 1.1921 | 12.8626 | [5 empty] | -0.024454
-24.2 | -0.0012067 | 1.1921 | 0 | 12.8291 | 1.1921 | 1.264 | 1.264 | 1.2619 | 1.1906 | 1e-05 | 1e-05 |  | 0.00011352 | -0.0012067 | -24.2 | 1.1906 | 12.8291 | [5 empty] | -0.027508
-24.1 | -0.0012067 | 1.1921 | -0.033601 | 12.7956 | 1.1921 | 1.264 | 1.264 | 1.2598 | 1.1891 | 1e-05 | 1e-05 |  | 0.00011738 | -0.0012067 | -24.1 | 1.1891 | 12.7956 | [5 empty] | -0.027492
-24 | -0.0011696 | 1.1888 | -0.033601 | 12.7621 | 1.1888 | 1.2593 | 1.2593 | 1.2577 | 1.1876 | 1e-05 | 0.022418 |  | 0.00012138 | -0.0011696 | -24 | 1.1876 | 12.7621 | [5 empty] | -0.024437
-23.9 | -0.0011696 | 1.1854 | -0.033601 | 12.7286 | 1.1854 | 1.2545 | 1.2545 | 1.2556 | 1.1862 | 1e-05 | 0.020945 |  | 0.00012552 | -0.0011696 | -23.9 | 1.1862 | 12.7286 | [5 empty] | -0.021383
-23.8 | -0.0012067 | 1.182 | -0.033601 | 12.6951 | 1.182 | 1.2498 | 1.2498 | 1.2536 | 1.1847 | 1e-05 | 1e-05 |  | 0.00012979 | -0.0012067 | -23.8 | 1.1847 | 12.6951 | [5 empty] | -0.024437
-23.7 | -0.0011696 | 1.1787 | -0.033601 | 12.6616 | 1.1787 | 1.2451 | 1.2451 | 1.2515 | 1.1832 | 1e-05 | 0.017996 |  | 0.00013421 | -0.0011696 | -23.7 | 1.1832 | 12.6616 | [5 empty] | -0.021383
-23.6 | -0.0012067 | 1.1753 | -0.033601 | 12.6281 | 1.1753 | 1.2403 | 1.2403 | 1.2494 | 1.1818 | 1e-05 | 1e-05 |  | 0.00013878 | -0.0012067 | -23.6 | 1.1818 | 12.6281 | [5 empty] | -0.024437
-23.5 | -0.0011696 | 1.172 | 0 | 12.5947 | 1.172 | 1.2356 | 1.2356 | 1.2473 | 1.1803 | 1e-05 | 0.015045 |  | 0.00014351 | -0.0011696 | -23.5 | 1.1803 | 12.5947 | [5 empty] | -0.021383
-23.4 | -0.0011696 | 1.172 | 0 | 12.5612 | 1.172 | 1.2356 | 1.2356 | 1.2452 | 1.1788 | 1e-05 | 0.013569 |  | 0.0001484 | -0.0011696 | -23.4 | 1.1788 | 12.5612 | [5 empty] | -0.018328
-23.3 | -0.0011696 | 1.172 | -0.033601 | 12.5277 | 1.172 | 1.2356 | 1.2356 | 1.2432 | 1.1773 | 1e-05 | 0.012092 |  | 0.00015346 | -0.0011696 | -23.3 | 1.1773 | 12.5277 | [5 empty] | -0.015273
-23.2 | -0.0011696 | 1.1686 | 0 | 12.4942 | 1.1686 | 1.2309 | 1.2309 | 1.2411 | 1.1758 | 1e-05 | 0.010614 |  | 0.00015868 | -0.0011696 | -23.2 | 1.1758 | 12.4942 | [5 empty] | -0.012219
-23.1 | -0.0011696 | 1.1686 | -0.033601 | 12.4607 | 1.1686 | 1.2309 | 1.2309 | 1.239 | 1.1744 | 1e-05 | 0.009136 |  | 0.00016409 | -0.0011696 | -23.1 | 1.1744 | 12.4607 | [5 empty] | -0.009164
-23 | -0.0011696 | 1.1652 | 0 | 12.4272 | 1.1652 | 1.2262 | 1.2262 | 1.2369 | 1.1729 | 1e-05 | 0.0076571 |  | 0.00016968 | -0.0011696 | -23 | 1.1729 | 12.4272 | [5 empty] | -0.0061093
-22.9 | -0.0011696 | 1.1652 | 0 | 12.3937 | 1.1652 | 1.2262 | 1.2262 | 1.2348 | 1.1714 | 1e-05 | 0.0061777 |  | 0.00017546 | -0.0011696 | -22.9 | 1.1714 | 12.3937 | [5 empty] | -0.0091722
-22.8 | -0.0011696 | 1.1652 | 0 | 12.3602 | 1.1652 | 1.2262 | 1.2262 | 1.2328 | 1.1699 | 1e-05 | 0.0046976 |  | 0.00018143 | -0.0011696 | -22.8 | 1.1699 | 12.3602 | [5 empty] | -0.0091722
-22.7 | -0.0011696 | 1.1652 | 0 | 12.3267 | 1.1652 | 1.2262 | 1.2262 | 1.2307 | 1.1684 | 1e-05 | 0.0032168 |  | 0.00018762 | -0.0011696 | -22.7 | 1.1684 | 12.3267 | [5 empty] | -0.0061176
-22.6 | -0.0011696 | 1.1652 | 0 | 12.2932 | 1.1652 | 1.2262 | 1.2262 | 1.2286 | 1.167 | 1e-05 | 0.0017355 |  | 0.00019401 | -0.0011696 | -22.6 | 1.167 | 12.2932 | [5 empty] | -0.0061176
-22.5 | -0.0011696 | 1.1652 | 0 | 12.2597 | 1.1652 | 1.2262 | 1.2262 | 1.2265 | 1.1655 | 1e-05 | 0.00025351 |  | 0.00020061 | -0.0011696 | -22.5 | 1.1655 | 12.2597 | [5 empty] | -0.0030629
-22.4 | -0.0011696 | 1.1652 | -0.033692 | 12.2262 | 1.1652 | 1.2262 | 1.2262 | 1.2244 | 1.164 | 1e-05 | 1e-05 |  | 0.00020745 | -0.0011696 | -22.4 | 1.164 | 12.2262 | [5 empty] | -0.0030629
-22.3 | -0.0011696 | 1.1619 | 0 | 12.1927 | 1.1619 | 1.2214 | 1.2214 | 1.2224 | 1.1625 | 1e-05 | 1e-05 |  | 0.00021451 | -0.0011696 | -22.3 | 1.1625 | 12.1927 | [5 empty] | -0.0061258
-22.2 | -0.0011696 | 1.1619 | 0 | 12.1592 | 1.1619 | 1.2214 | 1.2214 | 1.2203 | 1.161 | 1e-05 | 1e-05 |  | 0.00022182 | -0.0011696 | -22.2 | 1.161 | 12.1592 | [5 empty] | -0.0091887
-22.1 | -0.0011696 | 1.1619 | 0 | 12.1257 | 1.1619 | 1.2214 | 1.2214 | 1.2182 | 1.1595 | 1e-05 | 1e-05 |  | 0.00022938 | -0.0011696 | -22.1 | 1.1595 | 12.1257 | [5 empty] | -0.012252
-22 | -0.0011696 | 1.1619 | 0 | 12.0922 | 1.1619 | 1.2214 | 1.2214 | 1.2161 | 1.1581 | 1e-05 | 1e-05 |  | 0.00023719 | -0.0011696 | -22 | 1.1581 | 12.0922 | [5 empty] | -0.015314
-21.9 | -0.0011696 | 1.1619 | 0 | 12.0587 | 1.1619 | 1.2214 | 1.2214 | 1.214 | 1.1566 | 1e-05 | 1e-05 |  | 0.00024527 | -0.0011696 | -21.9 | 1.1566 | 12.0587 | [5 empty] | -0.018377
-21.8 | -0.0011324 | 1.1619 | -0.033692 | 12.0252 | 1.1619 | 1.2214 | 1.2214 | 1.212 | 1.1551 | 1e-05 | 0.026923 |  | 0.00025363 | -0.0011324 | -21.8 | 1.1551 | 12.0252 | [5 empty] | -0.015314
-21.7 | -0.0011696 | 1.1585 | -0.033692 | 11.9917 | 1.1585 | 1.2167 | 1.2167 | 1.2099 | 1.1536 | 1e-05 | 1e-05 |  | 0.00026227 | -0.0011696 | -21.7 | 1.1536 | 11.9917 | [5 empty] | -0.018377
-21.6 | -0.0011696 | 1.1551 | -0.033692 | 11.9582 | 1.1551 | 1.212 | 1.212 | 1.2078 | 1.1521 | 1e-05 | 1e-05 |  | 0.0002712 | -0.0011696 | -21.6 | 1.1521 | 11.9582 | [5 empty] | -0.02144
-21.5 | -0.0011696 | 1.1518 | -0.033692 | 11.9247 | 1.1518 | 1.2073 | 1.2073 | 1.2057 | 1.1506 | 1e-05 | 1e-05 |  | 0.00028044 | -0.0011696 | -21.5 | 1.1506 | 11.9247 | [5 empty] | -0.024503
-21.4 | -0.0011696 | 1.1484 | -0.033692 | 11.8912 | 1.1484 | 1.2026 | 1.2026 | 1.2036 | 1.1491 | 1e-05 | 1e-05 |  | 0.00028999 | -0.0011696 | -21.4 | 1.1491 | 11.8912 | [5 empty] | -0.027566
-21.3 | -0.0011696 | 1.145 | 0 | 11.8577 | 1.145 | 1.1979 | 1.1979 | 1.2016 | 1.1477 | 1e-05 | 1e-05 |  | 0.00029987 | -0.0011696 | -21.3 | 1.1477 | 11.8577 | [5 empty] | -0.030629
-21.2 | -0.0011324 | 1.145 | -0.033692 | 11.8242 | 1.145 | 1.1979 | 1.1979 | 1.1995 | 1.1462 | 1e-05 | 0.017991 |  | 0.00031008 | -0.0011324 | -21.2 | 1.1462 | 11.8242 | [5 empty] | -0.027566
-21.1 | -0.0011324 | 1.1416 | -0.033692 | 11.7907 | 1.1416 | 1.1932 | 1.1932 | 1.1974 | 1.1447 | 1e-05 | 0.0165 |  | 0.00032065 | -0.0011324 | -21.1 | 1.1447 | 11.7907 | [5 empty] | -0.024503
-21 | -0.0011324 | 1.1383 | -0.033692 | 11.7572 | 1.1383 | 1.1885 | 1.1885 | 1.1953 | 1.1432 | 1e-05 | 0.015009 |  | 0.00033157 | -0.0011324 | -21 | 1.1432 | 11.7572 | [5 empty] | -0.02144
-20.9 | -0.0011324 | 1.1349 | -0.033692 | 11.7237 | 1.1349 | 1.1838 | 1.1838 | 1.1932 | 1.1417 | 1e-05 | 0.013517 |  | 0.00034286 | -0.0011324 | -20.9 | 1.1417 | 11.7237 | [5 empty] | -0.018377
-20.8 | -0.0011324 | 1.1315 | -0.033692 | 11.6902 | 1.1315 | 1.1791 | 1.1791 | 1.1912 | 1.1402 | 1e-05 | 0.012024 |  | 0.00035454 | -0.0011324 | -20.8 | 1.1402 | 11.6902 | [5 empty] | -0.015314
-20.7 | -0.0011324 | 1.1282 | 0 | 11.6568 | 1.1282 | 1.1744 | 1.1744 | 1.1891 | 1.1387 | 1e-05 | 0.01053 |  | 0.00036662 | -0.0011324 | -20.7 | 1.1387 | 11.6568 | [5 empty] | -0.015314
-20.6 | -0.0011324 | 1.1282 | 0 | 11.6233 | 1.1282 | 1.1744 | 1.1744 | 1.187 | 1.1372 | 1e-05 | 0.009036 |  | 0.00037911 | -0.0011324 | -20.6 | 1.1372 | 11.6233 | [5 empty] | -0.012252
-20.5 | -0.0011324 | 1.1282 | 0 | 11.5898 | 1.1282 | 1.1744 | 1.1744 | 1.1849 | 1.1357 | 1e-05 | 0.0075412 |  | 0.00039202 | -0.0011324 | -20.5 | 1.1357 | 11.5898 | [5 empty] | -0.012252
-20.4 | -0.0011324 | 1.1282 | 0 | 11.5563 | 1.1282 | 1.1744 | 1.1744 | 1.1828 | 1.1342 | 1e-05 | 0.0060458 |  | 0.00040537 | -0.0011324 | -20.4 | 1.1342 | 11.5563 | [5 empty] | -0.0091887
-20.3 | -0.0011324 | 1.1282 | 0 | 11.5228 | 1.1282 | 1.1744 | 1.1744 | 1.1808 | 1.1327 | 1e-05 | 0.0045496 |  | 0.00041918 | -0.0011324 | -20.3 | 1.1327 | 11.5228 | [5 empty] | -0.0061258
-20.2 | -0.0011324 | 1.1282 | 0 | 11.4893 | 1.1282 | 1.1744 | 1.1744 | 1.1787 | 1.1312 | 1e-05 | 0.0030529 |  | 0.00043346 | -0.0011324 | -20.2 | 1.1312 | 11.4893 | [5 empty] | -0.0030629
-20.1 | -0.0011324 | 1.1282 | 0 | 11.4558 | 1.1282 | 1.1744 | 1.1744 | 1.1766 | 1.1297 | 1e-05 | 0.0015554 |  | 0.00044823 | -0.0011324 | -20.1 | 1.1297 | 11.4558 | [5 empty] | -0.0030629
-20 | -0.0011324 | 1.1282 | -0.033692 | 11.4223 | 1.1282 | 1.1744 | 1.1744 | 1.1745 | 1.1282 | 1e-05 | 5.7316e-05 |  | 0.00046349 | -0.0011324 | -20 | 1.1282 | 11.4223 | [5 empty] | -0.0030629
-19.9 | -0.0011324 | 1.1248 | 0 | 11.3888 | 1.1248 | 1.1698 | 1.1698 | 1.1724 | 1.1267 | 1e-05 | 1e-05 |  | 0.00047928 | -0.0011324 | -19.9 | 1.1267 | 11.3888 | [5 empty] | -0.0061258
-19.8 | -0.0011324 | 1.1248 | 0 | 11.3553 | 1.1248 | 1.1698 | 1.1698 | 1.1704 | 1.1252 | 1e-05 | 1e-05 |  | 0.00049561 | -0.0011324 | -19.8 | 1.1252 | 11.3553 | [5 empty] | -0.0091887
-19.7 | -0.0011324 | 1.1248 | 0 | 11.3218 | 1.1248 | 1.1698 | 1.1698 | 1.1683 | 1.1237 | 1e-05 | 1e-05 |  | 0.00051249 | -0.0011324 | -19.7 | 1.1237 | 11.3218 | [5 empty] | -0.0091887
-19.6 | -0.0011324 | 1.1248 | 0 | 11.2883 | 1.1248 | 1.1698 | 1.1698 | 1.1662 | 1.1222 | 1e-05 | 1e-05 |  | 0.00052995 | -0.0011324 | -19.6 | 1.1222 | 11.2883 | [5 empty] | -0.012252
-19.5 | -0.0011324 | 1.1248 | 0 | 11.2548 | 1.1248 | 1.1698 | 1.1698 | 1.1641 | 1.1207 | 1e-05 | 1e-05 |  | 0.000548 | -0.0011324 | -19.5 | 1.1207 | 11.2548 | [5 empty] | -0.015314
-19.4 | -0.0010952 | 1.1248 | -0.033692 | 11.2213 | 1.1248 | 1.1698 | 1.1698 | 1.162 | 1.1192 | 1e-05 | 0.028116 |  | 0.00056667 | -0.0010952 | -19.4 | 1.1192 | 11.2213 | [5 empty] | -0.012252
-19.3 | -0.0011324 | 1.1214 | -0.033692 | 11.1878 | 1.1214 | 1.1651 | 1.1651 | 1.16 | 1.1177 | 1e-05 | 1e-05 |  | 0.00058597 | -0.0011324 | -19.3 | 1.1177 | 11.1878 | [5 empty] | -0.015314
-19.2 | -0.0011324 | 1.1181 | 0 | 11.1543 | 1.1181 | 1.1604 | 1.1604 | 1.1579 | 1.1162 | 1e-05 | 1e-05 |  | 0.00060593 | -0.0011324 | -19.2 | 1.1162 | 11.1543 | [5 empty] | -0.018377
-19.1 | -0.0011324 | 1.1181 | -0.033692 | 11.1208 | 1.1181 | 1.1604 | 1.1604 | 1.1558 | 1.1147 | 1e-05 | 1e-05 |  | 0.00062657 | -0.0011324 | -19.1 | 1.1147 | 11.1208 | [5 empty] | -0.02144
-19 | -0.0011324 | 1.1147 | -0.033692 | 11.0873 | 1.1147 | 1.1558 | 1.1558 | 1.1537 | 1.1132 | 1e-05 | 1e-05 |  | 0.00064791 | -0.0011324 | -19 | 1.1132 | 11.0873 | [5 empty] | -0.024503
-18.9 | -0.0011324 | 1.1113 | 0 | 11.0538 | 1.1113 | 1.1511 | 1.1511 | 1.1516 | 1.1117 | 1e-05 | 1e-05 |  | 0.00066999 | -0.0011324 | -18.9 | 1.1117 | 11.0538 | [5 empty] | -0.027566
-18.8 | -0.0010952 | 1.1113 | -0.033692 | 11.0203 | 1.1113 | 1.1511 | 1.1511 | 1.1496 | 1.1102 | 1e-05 | 0.019089 |  | 0.00069281 | -0.0010952 | -18.8 | 1.1102 | 11.0203 | [5 empty] | -0.024503
-18.7 | -0.0010952 | 1.108 | -0.033692 | 10.9868 | 1.108 | 1.1465 | 1.1465 | 1.1475 | 1.1087 | 1e-05 | 0.017582 |  | 0.00071641 | -0.0010952 | -18.7 | 1.1087 | 10.9868 | [5 empty] | -0.02144
-18.6 | -0.0011324 | 1.1046 | -0.033692 | 10.9533 | 1.1046 | 1.1418 | 1.1418 | 1.1454 | 1.1072 | 1e-05 | 1e-05 |  | 0.00074081 | -0.0011324 | -18.6 | 1.1072 | 10.9533 | [5 empty] | -0.024503
-18.5 | -0.0010952 | 1.1012 | -0.033692 | 10.9198 | 1.1012 | 1.1372 | 1.1372 | 1.1433 | 1.1057 | 1e-05 | 0.014566 |  | 0.00076605 | -0.0010952 | -18.5 | 1.1057 | 10.9198 | [5 empty] | -0.02144
-18.4 | -0.0010952 | 1.0978 | -0.033692 | 10.8863 | 1.0978 | 1.1325 | 1.1325 | 1.1412 | 1.1042 | 1e-05 | 0.013058 |  | 0.00079214 | -0.0010952 | -18.4 | 1.1042 | 10.8863 | [5 empty] | -0.02144
-18.3 | -0.0010952 | 1.0945 | 0 | 10.8528 | 1.0945 | 1.1279 | 1.1279 | 1.1392 | 1.1027 | 1e-05 | 0.011548 |  | 0.00081912 | -0.0010952 | -18.3 | 1.1027 | 10.8528 | [5 empty] | -0.02144
-18.2 | -0.0010952 | 1.0945 | 0 | 10.8193 | 1.0945 | 1.1279 | 1.1279 | 1.1371 | 1.1011 | 1e-05 | 0.010038 |  | 0.00084703 | -0.0010952 | -18.2 | 1.1011 | 10.8193 | [5 empty] | -0.02144
-18.1 | -0.0010952 | 1.0945 | -0.033692 | 10.7858 | 1.0945 | 1.1279 | 1.1279 | 1.135 | 1.0996 | 1e-05 | 0.0085267 |  | 0.00087588 | -0.0010952 | -18.1 | 1.0996 | 10.7858 | [5 empty] | -0.02144
-18 | -0.0010952 | 1.0911 | 0 | 10.7523 | 1.0911 | 1.1233 | 1.1233 | 1.1329 | 1.0981 | 1e-05 | 0.007015 |  | 0.00090571 | -0.0010952 | -18 | 1.0981 | 10.7523 | [5 empty] | -0.02144
-17.9 | -0.0010952 | 1.0911 | -0.033692 | 10.7189 | 1.0911 | 1.1233 | 1.1233 | 1.1308 | 1.0966 | 1e-05 | 0.0055027 |  | 0.00093657 | -0.0010952 | -17.9 | 1.0966 | 10.7189 | [5 empty] | -0.02144
-17.8 | -0.0010952 | 1.0877 | 0 | 10.6854 | 1.0877 | 1.1187 | 1.1187 | 1.1288 | 1.0951 | 1e-05 | 0.0039896 |  | 0.00096847 | -0.0010952 | -17.8 | 1.0951 | 10.6854 | [5 empty] | -0.02144
-17.7 | -0.0010952 | 1.0877 | -0.033692 | 10.6519 | 1.0877 | 1.1187 | 1.1187 | 1.1267 | 1.0936 | 1e-05 | 0.0024759 |  | 0.0010015 | -0.0010952 | -17.7 | 1.0936 | 10.6519 | [5 empty] | -0.024503
-17.6 | -0.0010952 | 1.0844 | -0.033692 | 10.6184 | 1.0844 | 1.114 | 1.114 | 1.1246 | 1.0921 | 1e-05 | 0.00096146 |  | 0.0010356 | -0.0010952 | -17.6 | 1.0921 | 10.6184 | [5 empty] | -0.027566
-17.5 | -0.0010952 | 1.081 | -0.033692 | 10.5849 | 1.081 | 1.1094 | 1.1094 | 1.1225 | 1.0906 | 1e-05 | 1e-05 |  | 0.0010708 | -0.0010952 | -17.5 | 1.0906 | 10.5849 | [5 empty] | -0.024503
-17.4 | -0.0010952 | 1.0776 | -0.033692 | 10.5514 | 1.0776 | 1.1048 | 1.1048 | 1.1204 | 1.089 | 1e-05 | 1e-05 |  | 0.0011073 | -0.0010952 | -17.4 | 1.089 | 10.5514 | [5 empty] | -0.024503
-17.3 | -0.001058 | 1.0743 | -0.033692 | 10.5179 | 1.0743 | 1.1002 | 1.1002 | 1.1184 | 1.0875 | 1e-05 | 0.033475 |  | 0.001145 | -0.001058 | -17.3 | 1.0875 | 10.5179 | [5 empty] | -0.02144
-17.2 | -0.0010952 | 1.0709 | -0.033692 | 10.4844 | 1.0709 | 1.0956 | 1.0956 | 1.1163 | 1.086 | 1e-05 | 1e-05 |  | 0.001184 | -0.0010952 | -17.2 | 1.086 | 10.4844 | [5 empty] | -0.024503
-17.1 | -0.001058 | 1.0675 | -0.033692 | 10.4509 | 1.0675 | 1.091 | 1.091 | 1.1142 | 1.0845 | 1e-05 | 0.03044 |  | 0.0012244 | -0.001058 | -17.1 | 1.0845 | 10.4509 | [5 empty] | -0.02144
-17 | -0.001058 | 1.0642 | 0 | 10.4174 | 1.0642 | 1.0864 | 1.0864 | 1.1121 | 1.083 | 1e-05 | 0.028921 |  | 0.0012661 | -0.001058 | -17 | 1.083 | 10.4174 | [5 empty] | -0.018377
-16.9 | -0.001058 | 1.0642 | 0 | 10.3839 | 1.0642 | 1.0864 | 1.0864 | 1.11 | 1.0814 | 1e-05 | 0.027402 |  | 0.0013092 | -0.001058 | -16.9 | 1.0814 | 10.3839 | [5 empty] | -0.015314
-16.8 | -0.001058 | 1.0642 | 0 | 10.3504 | 1.0642 | 1.0864 | 1.0864 | 1.108 | 1.0799 | 1e-05 | 0.025882 |  | 0.0013538 | -0.001058 | -16.8 | 1.0799 | 10.3504 | [5 empty] | -0.015306
-16.7 | -0.001058 | 1.0642 | -0.033692 | 10.3169 | 1.0642 | 1.0864 | 1.0864 | 1.1059 | 1.0784 | 1e-05 | 0.024361 |  | 0.0013999 | -0.001058 | -16.7 | 1.0784 | 10.3169 | [5 empty] | -0.012243
-16.6 | -0.001058 | 1.0608 | 0 | 10.2834 | 1.0608 | 1.0819 | 1.0819 | 1.1038 | 1.0769 | 1e-05 | 0.022839 |  | 0.0014476 | -0.001058 | -16.6 | 1.0769 | 10.2834 | [5 empty] | -0.0091804
-16.5 | -0.001058 | 1.0608 | 0 | 10.2499 | 1.0608 | 1.0819 | 1.0819 | 1.1017 | 1.0754 | 1e-05 | 0.021317 |  | 0.0014969 | -0.001058 | -16.5 | 1.0754 | 10.2499 | [5 empty] | -0.0061176
-16.4 | -0.0010952 | 1.0608 | 0 | 10.2164 | 1.0608 | 1.0819 | 1.0819 | 1.0996 | 1.0738 | 1e-05 | 1e-05 |  | 0.0015479 | -0.0010952 | -16.4 | 1.0738 | 10.2164 | [5 empty] | -0.0091804
-16.3 | -0.0010952 | 1.0608 | -0.033601 | 10.1829 | 1.0608 | 1.0819 | 1.0819 | 1.0976 | 1.0723 | 1e-05 | 1e-05 |  | 0.0016006 | -0.0010952 | -16.3 | 1.0723 | 10.1829 | [5 empty] | -0.015298
-16.2 | -0.001058 | 1.0574 | 0 | 10.1494 | 1.0574 | 1.0773 | 1.0773 | 1.0955 | 1.0708 | 1e-05 | 0.016746 |  | 0.0016552 | -0.001058 | -16.2 | 1.0708 | 10.1494 | [5 empty] | -0.015298
-16.1 | -0.001058 | 1.0574 | 0 | 10.1159 | 1.0574 | 1.0773 | 1.0773 | 1.0934 | 1.0693 | 1e-05 | 0.015221 |  | 0.0017116 | -0.001058 | -16.1 | 1.0693 | 10.1159 | [5 empty] | -0.012235
-16 | -0.001058 | 1.0574 | 0 | 10.0824 | 1.0574 | 1.0773 | 1.0773 | 1.0913 | 1.0677 | 1e-05 | 0.013695 |  | 0.0017699 | -0.001058 | -16 | 1.0677 | 10.0824 | [5 empty] | -0.012235
-15.9 | -0.001058 | 1.0574 | -0.033692 | 10.0489 | 1.0574 | 1.0773 | 1.0773 | 1.0892 | 1.0662 | 1e-05 | 0.012169 |  | 0.0018301 | -0.001058 | -15.9 | 1.0662 | 10.0489 | [5 empty] | -0.012235
-15.8 | -0.001058 | 1.0541 | -0.067293 | 10.0154 | 1.0541 | 1.0727 | 1.0727 | 1.0872 | 1.0647 | 1e-05 | 0.010641 |  | 0.0018925 | -0.001058 | -15.8 | 1.0647 | 10.0154 | [5 empty] | -0.012235
-15.7 | -0.0010209 | 1.0473 | 0 | 9.9819 | 1.0473 | 1.0636 | 1.0636 | 1.0851 | 1.0632 | 1e-05 | 0.046075 |  | 0.0019569 | -0.0010209 | -15.7 | 1.0632 | 9.9819 | [5 empty] | -0.0091804
-15.6 | -0.001058 | 1.0473 | 0 | 9.9484 | 1.0473 | 1.0636 | 1.0636 | 1.083 | 1.0616 | 1e-05 | 0.0075844 |  | 0.0020236 | -0.001058 | -15.6 | 1.0616 | 9.9484 | [5 empty] | -0.012235
-15.5 | -0.001058 | 1.0473 | 0 | 9.9149 | 1.0473 | 1.0636 | 1.0636 | 1.0809 | 1.0601 | 1e-05 | 0.0060549 |  | 0.0020925 | -0.001058 | -15.5 | 1.0601 | 9.9149 | [5 empty] | -0.01529
-15.4 | -0.001058 | 1.0473 | 0 | 9.8814 | 1.0473 | 1.0636 | 1.0636 | 1.0788 | 1.0586 | 1e-05 | 0.0045247 |  | 0.0021638 | -0.001058 | -15.4 | 1.0586 | 9.8814 | [5 empty] | -0.018344
-15.3 | -0.001058 | 1.0473 | 0 | 9.8479 | 1.0473 | 1.0636 | 1.0636 | 1.0768 | 1.057 | 1e-05 | 0.0029937 |  | 0.0022375 | -0.001058 | -15.3 | 1.057 | 9.8479 | [5 empty] | -0.018336
-15.2 | -0.0010209 | 1.0473 | 0 | 9.8144 | 1.0473 | 1.0636 | 1.0636 | 1.0747 | 1.0555 | 1e-05 | 0.038423 |  | 0.0023137 | -0.0010209 | -15.2 | 1.0555 | 9.8144 | [5 empty] | -0.009164
-15.1 | -0.001058 | 1.0473 | -0.033601 | 9.781 | 1.0473 | 1.0636 | 1.0636 | 1.0726 | 1.054 | 1e-05 | 1e-05 |  | 0.0023926 | -0.001058 | -15.1 | 1.054 | 9.781 | [5 empty] | -0.012219
-15 | -0.001058 | 1.044 | -0.033601 | 9.7475 | 1.044 | 1.059 | 1.059 | 1.0705 | 1.0524 | 1e-05 | 1e-05 |  | 0.0024741 | -0.001058 | -15 | 1.0524 | 9.7475 | [5 empty] | -0.015273
-14.9 | -0.001058 | 1.0406 | -0.033601 | 9.714 | 1.0406 | 1.0545 | 1.0545 | 1.0684 | 1.0509 | 1e-05 | 1e-05 |  | 0.0025583 | -0.001058 | -14.9 | 1.0509 | 9.714 | [5 empty] | -0.018328
-14.8 | -0.001058 | 1.0372 | -0.033601 | 9.6805 | 1.0372 | 1.05 | 1.05 | 1.0664 | 1.0494 | 1e-05 | 1e-05 |  | 0.0026455 | -0.001058 | -14.8 | 1.0494 | 9.6805 | [5 empty] | -0.021383
-14.7 | -0.001058 | 1.0339 | 0.033601 | 9.647 | 1.0339 | 1.0454 | 1.0454 | 1.0643 | 1.0478 | 1e-05 | 1e-05 |  | 0.0027356 | -0.001058 | -14.7 | 1.0478 | 9.647 | [5 empty] | -0.024437
-14.6 | -0.0010209 | 1.0372 | -0.033601 | 9.6135 | 1.0372 | 1.05 | 1.05 | 1.0622 | 1.0463 | 1e-05 | 0.029218 |  | 0.0028288 | -0.0010209 | -14.6 | 1.0463 | 9.6135 | [5 empty] | -0.024437
-14.5 | -0.0010209 | 1.0339 | -0.033601 | 9.58 | 1.0339 | 1.0454 | 1.0454 | 1.0601 | 1.0448 | 1e-05 | 0.027681 |  | 0.0029251 | -0.0010209 | -14.5 | 1.0448 | 9.58 | [5 empty] | -0.021383
-14.4 | -0.0010209 | 1.0305 | -0.033601 | 9.5465 | 1.0305 | 1.0409 | 1.0409 | 1.058 | 1.0432 | 1e-05 | 0.026143 |  | 0.0030248 | -0.0010209 | -14.4 | 1.0432 | 9.5465 | [5 empty] | -0.021391
-14.3 | -0.0010209 | 1.0272 | -0.033601 | 9.513 | 1.0272 | 1.0364 | 1.0364 | 1.056 | 1.0417 | 1e-05 | 0.024604 |  | 0.0031278 | -0.0010209 | -14.3 | 1.0417 | 9.513 | [5 empty] | -0.018336
-14.2 | -0.0010209 | 1.0238 | -0.033601 | 9.4795 | 1.0238 | 1.0319 | 1.0319 | 1.0539 | 1.0402 | 1e-05 | 0.023065 |  | 0.0032344 | -0.0010209 | -14.2 | 1.0402 | 9.4795 | [5 empty] | -0.015282
-14.1 | -0.001058 | 1.0204 | 0 | 9.446 | 1.0204 | 1.0274 | 1.0274 | 1.0518 | 1.0386 | 1e-05 | 1e-05 |  | 0.0033445 | -0.001058 | -14.1 | 1.0386 | 9.446 | [5 empty] | -0.024454
-14 | -0.0010209 | 1.0204 | 0 | 9.4125 | 1.0204 | 1.0274 | 1.0274 | 1.0497 | 1.0371 | 1e-05 | 0.019985 |  | 0.0034585 | -0.0010209 | -14 | 1.0371 | 9.4125 | [5 empty] | -0.024462
-13.9 | -0.0010209 | 1.0204 | -0.033692 | 9.379 | 1.0204 | 1.0274 | 1.0274 | 1.0477 | 1.0355 | 1e-05 | 0.018443 |  | 0.0035763 | -0.0010209 | -13.9 | 1.0355 | 9.379 | [5 empty] | -0.02447
-13.8 | -0.0010209 | 1.0171 | 0 | 9.3455 | 1.0171 | 1.0228 | 1.0228 | 1.0456 | 1.034 | 1e-05 | 0.016901 |  | 0.0036981 | -0.0010209 | -13.8 | 1.034 | 9.3455 | [5 empty] | -0.021416
-13.7 | -0.0010209 | 1.0171 | 0 | 9.312 | 1.0171 | 1.0228 | 1.0228 | 1.0435 | 1.0324 | 1e-05 | 0.015358 |  | 0.0038241 | -0.0010209 | -13.7 | 1.0324 | 9.312 | [5 empty] | -0.018361
-13.6 | -0.0010209 | 1.0171 | -0.067293 | 9.2785 | 1.0171 | 1.0228 | 1.0228 | 1.0414 | 1.0309 | 1e-05 | 0.013814 |  | 0.0039543 | -0.0010209 | -13.6 | 1.0309 | 9.2785 | [5 empty] | -0.018369
-13.5 | -0.0010209 | 1.0103 | -0.033692 | 9.245 | 1.0103 | 1.0138 | 1.0138 | 1.0393 | 1.0294 | 1e-05 | 0.012269 |  | 0.004089 | -0.0010209 | -13.5 | 1.0294 | 9.245 | [5 empty] | -0.021432
-13.4 | -0.0010209 | 1.007 | -0.033692 | 9.2115 | 1.007 | 1.0093 | 1.0093 | 1.0373 | 1.0278 | 1e-05 | 0.010724 |  | 0.0042283 | -0.0010209 | -13.4 | 1.0278 | 9.2115 | [5 empty] | -0.024495
-13.3 | -0.0009837 | 1.0036 | 0 | 9.178 | 1.0036 | 1.0048 | 1.0048 | 1.0352 | 1.0263 | 1e-05 | 0.046239 |  | 0.0043723 | -0.0009837 | -13.3 | 1.0263 | 9.178 | [5 empty] | -0.021432
-13.2 | -0.0010209 | 1.0036 | 0 | 9.1445 | 1.0036 | 1.0048 | 1.0048 | 1.0331 | 1.0247 | 1e-05 | 0.0076311 |  | 0.0045213 | -0.0010209 | -13.2 | 1.0247 | 9.1445 | [5 empty] | -0.024495
-13.1 | -0.0010209 | 1.0036 | -0.033692 | 9.111 | 1.0036 | 1.0048 | 1.0048 | 1.031 | 1.0232 | 1e-05 | 0.0060834 |  | 0.0046753 | -0.0010209 | -13.1 | 1.0232 | 9.111 | [5 empty] | -0.027558
-13 | -0.0009837 | 1.0002 | -0.033692 | 9.0775 | 1.0002 | 1.0003 | 1.0003 | 1.0289 | 1.0216 | 1e-05 | 0.041596 |  | 0.0048346 | -0.0009837 | -13 | 1.0216 | 9.0775 | [5 empty] | -0.02144
-12.9 | -0.0009837 | 0.99687 | -0.033692 | 9.044 | 0.99687 | 0.99583 | 0.99583 | 1.0269 | 1.0201 | 1e-05 | 0.040047 |  | 0.0049992 | -0.0009837 | -12.9 | 1.0201 | 9.044 | [5 empty] | -0.018377
-12.8 | -0.0009837 | 0.9935 | 0 | 9.0105 | 0.9935 | 0.99134 | 0.99134 | 1.0248 | 1.0185 | 1e-05 | 0.038497 |  | 0.0051695 | -0.0009837 | -12.8 | 1.0185 | 9.0105 | [5 empty] | -0.015314
-12.7 | -0.0010209 | 0.9935 | -0.033692 | 8.977 | 0.9935 | 0.99134 | 0.99134 | 1.0227 | 1.017 | 1e-05 | 1e-05 |  | 0.0053456 | -0.0010209 | -12.7 | 1.017 | 8.977 | [5 empty] | -0.018377
-12.6 | -0.0010209 | 0.99013 | -0.033692 | 8.9435 | 0.99013 | 0.98686 | 0.98686 | 1.0206 | 1.0154 | 1e-05 | 1e-05 |  | 0.0055277 | -0.0010209 | -12.6 | 1.0154 | 8.9435 | [5 empty] | -0.024503
-12.5 | -0.0009837 | 0.98676 | 0 | 8.91 | 0.98676 | 0.98239 | 0.98239 | 1.0185 | 1.0139 | 1e-05 | 0.033842 |  | 0.005716 | -0.0009837 | -12.5 | 1.0139 | 8.91 | [5 empty] | -0.024503
-12.4 | -0.0009837 | 0.98676 | 0 | 8.8766 | 0.98676 | 0.98239 | 0.98239 | 1.0165 | 1.0123 | 1e-05 | 0.032289 |  | 0.0059107 | -0.0009837 | -12.4 | 1.0123 | 8.8766 | [5 empty] | -0.024503
-12.3 | -0.0009837 | 0.98676 | 0 | 8.8431 | 0.98676 | 0.98239 | 0.98239 | 1.0144 | 1.0108 | 1e-05 | 0.030735 |  | 0.0061121 | -0.0009837 | -12.3 | 1.0108 | 8.8431 | [5 empty] | -0.02144
-12.2 | -0.0009837 | 0.98676 | -0.033692 | 8.8096 | 0.98676 | 0.98239 | 0.98239 | 1.0123 | 1.0092 | 1e-05 | 0.02918 |  | 0.0063203 | -0.0009837 | -12.2 | 1.0092 | 8.8096 | [5 empty] | -0.02144
-12.1 | -0.0009837 | 0.98339 | -0.067384 | 8.7761 | 0.98339 | 0.97792 | 0.97792 | 1.0102 | 1.0076 | 1e-05 | 0.027625 |  | 0.0065356 | -0.0009837 | -12.1 | 1.0076 | 8.7761 | [5 empty] | -0.018377
-12 | -0.0009837 | 0.97666 | -0.033692 | 8.7426 | 0.97666 | 0.969 | 0.969 | 1.0081 | 1.0061 | 1e-05 | 0.026068 |  | 0.0067582 | -0.0009837 | -12 | 1.0061 | 8.7426 | [5 empty] | -0.015314
-11.9 | -0.0009837 | 0.97329 | -0.033692 | 8.7091 | 0.97329 | 0.96454 | 0.96454 | 1.0061 | 1.0045 | 1e-05 | 0.024511 |  | 0.0069884 | -0.0009837 | -11.9 | 1.0045 | 8.7091 | [5 empty] | -0.018377
-11.8 | -0.0009837 | 0.96992 | 0 | 8.6756 | 0.96992 | 0.96009 | 0.96009 | 1.004 | 1.003 | 1e-05 | 0.022953 |  | 0.0072264 | -0.0009837 | -11.8 | 1.003 | 8.6756 | [5 empty] | -0.02144
-11.7 | -0.0009837 | 0.96992 | 0 | 8.6421 | 0.96992 | 0.96009 | 0.96009 | 1.0019 | 1.0014 | 1e-05 | 0.021394 |  | 0.0074726 | -0.0009837 | -11.7 | 1.0014 | 8.6421 | [5 empty] | -0.02144
-11.6 | -0.0009837 | 0.96992 | 0 | 8.6086 | 0.96992 | 0.96009 | 0.96009 | 0.99981 | 0.99986 | 1e-05 | 0.019835 |  | 0.0077271 | -0.0009837 | -11.6 | 0.99986 | 8.6086 | [5 empty] | -0.02144
-11.5 | -0.0009465 | 0.96992 | 0 | 8.5751 | 0.96992 | 0.96009 | 0.96009 | 0.99773 | 0.9983 | 1e-05 | 0.055335 |  | 0.0079904 | -0.0009465 | -11.5 | 0.9983 | 8.5751 | [5 empty] | -0.012252
-11.4 | -0.0009465 | 0.96992 | -0.033692 | 8.5416 | 0.96992 | 0.96009 | 0.96009 | 0.99565 | 0.99674 | 1e-05 | 0.053774 |  | 0.0082625 | -0.0009465 | -11.4 | 0.99674 | 8.5416 | [5 empty] | -0.0091887
-11.3 | -0.0009465 | 0.96655 | -0.033692 | 8.5081 | 0.96655 | 0.95565 | 0.95565 | 0.99357 | 0.99518 | 1e-05 | 0.052212 |  | 0.008544 | -0.0009465 | -11.3 | 0.99518 | 8.5081 | [5 empty] | -0.0061258
-11.2 | -0.0009837 | 0.96318 | 0 | 8.4746 | 0.96318 | 0.95121 | 0.95121 | 0.99149 | 0.99361 | 1e-05 | 0.013588 |  | 0.008835 | -0.0009837 | -11.2 | 0.99361 | 8.4746 | [5 empty] | -0.0091887
-11.1 | -0.0009837 | 0.96318 | -0.033692 | 8.4411 | 0.96318 | 0.95121 | 0.95121 | 0.98941 | 0.99205 | 1e-05 | 0.012024 |  | 0.009136 | -0.0009837 | -11.1 | 0.99205 | 8.4411 | [5 empty] | -0.012252
-11 | -0.0009837 | 0.95981 | 0.033692 | 8.4076 | 0.95981 | 0.94678 | 0.94678 | 0.98733 | 0.99048 | 1e-05 | 0.01046 |  | 0.0094472 | -0.0009837 | -11 | 0.99048 | 8.4076 | [5 empty] | -0.015314
-10.9 | -0.0009837 | 0.96318 | 0 | 8.3741 | 0.96318 | 0.95121 | 0.95121 | 0.98525 | 0.98892 | 1e-05 | 0.0088947 |  | 0.009769 | -0.0009837 | -10.9 | 0.98892 | 8.3741 | [5 empty] | -0.021432
-10.8 | -0.0009465 | 0.96318 | 0 | 8.3406 | 0.96318 | 0.95121 | 0.95121 | 0.98317 | 0.98735 | 1e-05 | 0.04439 |  | 0.010102 | -0.0009465 | -10.8 | 0.98735 | 8.3406 | [5 empty] | -0.018369
-10.7 | -0.0009465 | 0.96318 | -0.033692 | 8.3071 | 0.96318 | 0.95121 | 0.95121 | 0.98109 | 0.98579 | 1e-05 | 0.042823 |  | 0.010446 | -0.0009465 | -10.7 | 0.98579 | 8.3071 | [5 empty] | -0.015306
-10.6 | -0.0009837 | 0.95981 | -0.033692 | 8.2736 | 0.95981 | 0.94678 | 0.94678 | 0.97901 | 0.98422 | 1e-05 | 0.0041938 |  | 0.010802 | -0.0009837 | -10.6 | 0.98422 | 8.2736 | [5 empty] | -0.018369
-10.5 | -0.0009465 | 0.95644 | -0.033692 | 8.2401 | 0.95644 | 0.94235 | 0.94235 | 0.97693 | 0.98265 | 1e-05 | 0.039686 |  | 0.01117 | -0.0009465 | -10.5 | 0.98265 | 8.2401 | [5 empty] | -0.018361
-10.4 | -0.0009837 | 0.95307 | -0.067293 | 8.2066 | 0.95307 | 0.93792 | 0.93792 | 0.97485 | 0.98108 | 1e-05 | 0.0010557 |  | 0.01155 | -0.0009837 | -10.4 | 0.98108 | 8.2066 | [5 empty] | -0.027541
-10.3 | -0.0009465 | 0.94634 | 0 | 8.1731 | 0.94634 | 0.9291 | 0.9291 | 0.97277 | 0.97951 | 1e-05 | 0.036546 |  | 0.011944 | -0.0009465 | -10.3 | 0.97951 | 8.1731 | [5 empty] | -0.027541
-10.2 | -0.0009465 | 0.94634 | 0 | 8.1396 | 0.94634 | 0.9291 | 0.9291 | 0.97069 | 0.97794 | 1e-05 | 0.034975 |  | 0.01235 | -0.0009465 | -10.2 | 0.97794 | 8.1396 | [5 empty] | -0.027541
-10.1 | -0.0009465 | 0.94634 | -0.033692 | 8.1061 | 0.94634 | 0.9291 | 0.9291 | 0.96861 | 0.97637 | 1e-05 | 0.033403 |  | 0.012771 | -0.0009465 | -10.1 | 0.97637 | 8.1061 | [5 empty] | -0.024478
-10 | -0.0009465 | 0.94297 | -0.033601 | 8.0726 | 0.94297 | 0.9247 | 0.9247 | 0.96653 | 0.97479 | 1e-05 | 0.03183 |  | 0.013206 | -0.0009465 | -10 | 0.97479 | 8.0726 | [5 empty] | -0.02447
-9.9 | -0.0009465 | 0.93961 | -0.067293 | 8.0391 | 0.93961 | 0.92031 | 0.92031 | 0.96445 | 0.97322 | 1e-05 | 0.030257 |  | 0.013656 | -0.0009465 | -9.9 | 0.97322 | 8.0391 | [5 empty] | -0.021407
-9.8 | -0.0009094 | 0.93288 | 0 | 8.0056 | 0.93288 | 0.91153 | 0.91153 | 0.96237 | 0.97165 | 1e-05 | 0.065644 |  | 0.014121 | -0.0009094 | -9.8 | 0.97165 | 8.0056 | [5 empty] | -0.012235
-9.7 | -0.0009465 | 0.93288 | 0 | 7.9721 | 0.93288 | 0.91153 | 0.91153 | 0.96029 | 0.97007 | 1e-05 | 0.027107 |  | 0.014602 | -0.0009465 | -9.7 | 0.97007 | 7.9721 | [5 empty] | -0.01529
-9.6 | -0.0009465 | 0.93288 | 0 | 7.9387 | 0.93288 | 0.91153 |  | 0.95821 | 0.96849 | 1e-05 | 0.025531 |  | 0.0151 | -0.0009465 | -9.6 | 0.96849 | 7.9387 | [5 empty] | -0.018344
-9.5 | -0.0009465 | 0.93288 | -0.033601 | 7.9052 | 0.93288 | 0.91153 |  | 0.95613 | 0.96692 | 1e-05 | 0.023954 |  | 0.015614 | -0.0009465 | -9.5 | 0.96692 | 7.9052 | [5 empty] | -0.018336
-9.4 | -0.0009094 | 0.92952 | 0 | 7.8717 | 0.92952 | 0.90715 |  | 0.95406 | 0.96534 | 1e-05 | 0.059337 |  | 0.016146 | -0.0009094 | -9.4 | 0.96534 | 7.8717 | [5 empty] | -0.015282
-9.3 | -0.0009094 | 0.92952 | 0.033601 | 7.8382 | 0.92952 | 0.90715 |  | 0.95198 | 0.96376 | 1e-05 | 0.057758 |  | 0.016696 | -0.0009094 | -9.3 | 0.96376 | 7.8382 | [5 empty] | -0.009164
-9.2 | -0.0009465 | 0.93288 | -0.033601 | 7.8047 | 0.93288 | 0.91153 |  | 0.9499 | 0.96218 | 1e-05 | 0.019217 |  | 0.017265 | -0.0009465 | -9.2 | 0.96218 | 7.8047 | [5 empty] | -0.012219
-9.1 | -0.0009465 | 0.92952 | -0.033601 | 7.7712 | 0.92952 | 0.90715 |  | 0.94782 | 0.9606 | 1e-05 | 0.017637 |  | 0.017853 | -0.0009465 | -9.1 | 0.9606 | 7.7712 | [5 empty] | -0.015273
-9 | -0.0009465 | 0.92616 | -0.033601 | 7.7377 | 0.92616 | 0.90278 |  | 0.94574 | 0.95902 | 1e-05 | 0.016055 |  | 0.018461 | -0.0009465 | -9 | 0.95902 | 7.7377 | [5 empty] | -0.018328
-8.9 | -0.0009094 | 0.9228 | 0 | 7.7042 | 0.9228 | 0.89842 |  | 0.94366 | 0.95744 | 1e-05 | 0.051434 |  | 0.01909 | -0.0009094 | -8.9 | 0.95744 | 7.7042 | [5 empty] | -0.015273
-8.8 | -0.0009465 | 0.9228 | 0 | 7.6707 | 0.9228 | 0.89842 |  | 0.94158 | 0.95585 | 1e-05 | 0.01289 |  | 0.01974 | -0.0009465 | -8.8 | 0.95585 | 7.6707 | [5 empty] | -0.018328
-8.7 | -0.0009465 | 0.9228 | -0.033601 | 7.6372 | 0.9228 | 0.89842 |  | 0.9395 | 0.95427 | 1e-05 | 0.011306 |  | 0.020412 | -0.0009465 | -8.7 | 0.95427 | 7.6372 | [5 empty] | -0.024437
-8.6 | -0.0009094 | 0.91944 | -0.033601 | 7.6037 | 0.91944 | 0.89406 |  | 0.93742 | 0.95268 | 1e-05 | 0.046682 |  | 0.021108 | -0.0009094 | -8.6 | 0.95268 | 7.6037 | [5 empty] | -0.021383
-8.5 | -0.0009094 | 0.91608 | -0.033601 | 7.5702 | 0.91608 | 0.88971 |  | 0.93534 | 0.9511 | 1e-05 | 0.045097 |  | 0.021827 | -0.0009094 | -8.5 | 0.9511 | 7.5702 | [5 empty] | -0.018328
-8.4 | -0.0009094 | 0.91272 | 0 | 7.5367 | 0.91272 | 0.88536 |  | 0.93326 | 0.94951 | 1e-05 | 0.04351 |  | 0.02257 | -0.0009094 | -8.4 | 0.94951 | 7.5367 | [5 empty] | -0.018344
-8.3 | -0.0009094 | 0.91272 | -0.033601 | 7.5032 | 0.91272 | 0.88536 |  | 0.93118 | 0.94792 | 1e-05 | 0.041922 |  | 0.023339 | -0.0009094 | -8.3 | 0.94792 | 7.5032 | [5 empty] | -0.018344
-8.2 | -0.0009094 | 0.90936 | -0.033601 | 7.4697 | 0.90936 | 0.88101 |  | 0.9291 | 0.94634 | 1e-05 | 0.040334 |  | 0.024134 | -0.0009094 | -8.2 | 0.94634 | 7.4697 | [5 empty] | -0.018344
-8.1 | -0.0009094 | 0.906 | 0 | 7.4362 | 0.906 | 0.87668 |  | 0.92702 | 0.94475 | 1e-05 | 0.038745 |  | 0.024956 | -0.0009094 | -8.1 | 0.94475 | 7.4362 | [5 empty] | -0.018361
-8 | -0.0009094 | 0.906 | 0 | 7.4027 | 0.906 | 0.87668 |  | 0.92494 | 0.94316 | 1e-05 | 0.037154 |  | 0.025806 | -0.0009094 | -8 | 0.94316 | 7.4027 | [5 empty] | -0.015306
-7.9 | -0.0009094 | 0.906 | -0.033782 | 7.3692 | 0.906 | 0.87668 |  | 0.92286 | 0.94157 | 1e-05 | 0.035563 |  | 0.026685 | -0.0009094 | -7.9 | 0.94157 | 7.3692 | [5 empty] | -0.015323
-7.8 | -0.0009094 | 0.90262 | 0 | 7.3357 | 0.90262 | 0.87232 |  | 0.92078 | 0.93997 | 1e-05 | 0.033971 |  | 0.027594 | -0.0009094 | -7.8 | 0.93997 | 7.3357 | [5 empty] | -0.018394
-7.7 | -0.0009094 | 0.90262 | 0 | 7.3022 | 0.90262 | 0.87232 |  | 0.9187 | 0.93838 | 1e-05 | 0.032378 |  | 0.028534 | -0.0009094 | -7.7 | 0.93838 | 7.3022 | [5 empty] | -0.01841
-7.6 | -0.0009094 | 0.90262 | -0.033782 | 7.2687 | 0.90262 | 0.87232 |  | 0.91662 | 0.93679 | 1e-05 | 0.030785 |  | 0.029506 | -0.0009094 | -7.6 | 0.93679 | 7.2687 | [5 empty] | -0.018427
-7.5 | -0.0009094 | 0.89925 | 0 | 7.2352 | 0.89925 | 0.86797 |  | 0.91454 | 0.93519 | 1e-05 | 0.02919 |  | 0.030511 | -0.0009094 | -7.5 | 0.93519 | 7.2352 | [5 empty] | -0.021498
-7.4 | -0.0009094 | 0.89925 | -0.033782 | 7.2017 | 0.89925 | 0.86797 |  | 0.91246 | 0.9336 | 1e-05 | 0.027594 |  | 0.03155 | -0.0009094 | -7.4 | 0.9336 | 7.2017 | [5 empty] | -0.024569
-7.3 | -0.0008721 | 0.89587 | -0.033782 | 7.1682 | 0.89587 | 0.86363 |  | 0.91038 | 0.932 | 1e-05 | 0.063159 |  | 0.032625 | -0.0008721 | -7.3 | 0.932 | 7.1682 | [5 empty] | -0.021498
-7.2 | -0.0009094 | 0.89249 | -0.033782 | 7.1347 | 0.89249 | 0.85929 |  | 0.9083 | 0.9304 | 1e-05 | 0.0244 |  | 0.033737 | -0.0009094 | -7.2 | 0.9304 | 7.1347 | [5 empty] | -0.024569
-7.1 | -0.0009094 | 0.88911 | -0.033782 | 7.1012 | 0.88911 | 0.85495 |  | 0.90622 | 0.9288 | 1e-05 | 0.022802 |  | 0.034886 | -0.0009094 | -7.1 | 0.9288 | 7.1012 | [5 empty] | -0.02764
-7 | -0.0008721 | 0.88573 | -0.033782 | 7.0677 | 0.88573 | 0.85062 |  | 0.90414 | 0.92721 | 1e-05 | 0.058363 |  | 0.036074 | -0.0008721 | -7 | 0.92721 | 7.0677 | [5 empty] | -0.027624
-6.9 | -0.0009094 | 0.88235 | -0.033782 | 7.0342 | 0.88235 | 0.8463 |  | 0.90206 | 0.92561 | 1e-05 | 0.019603 |  | 0.037303 | -0.0009094 | -6.9 | 0.92561 | 7.0342 | [5 empty] | -0.030695
-6.8 | -0.0008721 | 0.87898 | 0 | 7.0008 | 0.87898 | 0.84198 |  | 0.89998 | 0.924 | 1e-05 | 0.055162 |  | 0.038574 | -0.0008721 | -6.8 | 0.924 | 7.0008 | [5 empty] | -0.027624
-6.7 | -0.0008721 | 0.87898 | -0.033782 | 6.9673 | 0.87898 | 0.84198 |  | 0.8979 | 0.9224 | 1e-05 | 0.05356 |  | 0.039888 | -0.0008721 | -6.7 | 0.9224 | 6.9673 | [5 empty] | -0.027607
-6.6 | -0.0008721 | 0.8756 | -0.033782 | 6.9338 | 0.8756 | 0.83767 |  | 0.89582 | 0.9208 | 1e-05 | 0.051957 |  | 0.041246 | -0.0008721 | -6.6 | 0.9208 | 6.9338 | [5 empty] | -0.027591
-6.5 | -0.0008721 | 0.87222 | -0.033601 | 6.9003 | 0.87222 | 0.83336 |  | 0.89374 | 0.9192 | 1e-05 | 0.050353 |  | 0.042651 | -0.0008721 | -6.5 | 0.9192 | 6.9003 | [5 empty] | -0.027574
-6.4 | -0.0008721 | 0.86886 | -0.033782 | 6.8668 | 0.86886 | 0.82909 |  | 0.89166 | 0.91759 | 1e-05 | 0.048749 |  | 0.044104 | -0.0008721 | -6.4 | 0.91759 | 6.8668 | [5 empty] | -0.027558
-6.3 | -0.0008721 | 0.86548 | 0 | 6.8333 | 0.86548 | 0.82479 |  | 0.88958 | 0.91598 | 1e-05 | 0.047143 |  | 0.045606 | -0.0008721 | -6.3 | 0.91598 | 6.8333 | [5 empty] | -0.027541
-6.2 | -0.0008721 | 0.86548 | -0.033601 | 6.7998 | 0.86548 | 0.82479 |  | 0.8875 | 0.91438 | 1e-05 | 0.045536 |  | 0.04716 | -0.0008721 | -6.2 | 0.91438 | 6.7998 | [5 empty] | -0.030596
-6.1 | -0.0008721 | 0.86212 | -0.033601 | 6.7663 | 0.86212 | 0.82052 |  | 0.88542 | 0.91277 | 1e-05 | 0.043928 |  | 0.048766 | -0.0008721 | -6.1 | 0.91277 | 6.7663 | [5 empty] | -0.03058
-6 | -0.0008721 | 0.85876 | -0.033601 | 6.7328 | 0.85876 | 0.81626 |  | 0.88334 | 0.91116 | 1e-05 | 0.04232 |  | 0.050428 | -0.0008721 | -6 | 0.91116 | 6.7328 | [5 empty] | -0.030563
-5.9 | -0.000835 | 0.8554 | -0.033601 | 6.6993 | 0.8554 | 0.81201 |  | 0.88126 | 0.90955 | 1e-05 | 0.077672 |  | 0.052145 | -0.000835 | -5.9 | 0.90955 | 6.6993 | [5 empty] | -0.027508
-5.8 | -0.0008721 | 0.85204 | -0.033601 | 6.6658 | 0.85204 | 0.80776 |  | 0.87918 | 0.90794 | 1e-05 | 0.0391 |  | 0.053922 | -0.0008721 | -5.8 | 0.90794 | 6.6658 | [5 empty] | -0.030547
-5.7 | -0.0008721 | 0.84868 | -0.033601 | 6.6323 | 0.84868 | 0.80351 |  | 0.8771 | 0.90633 | 1e-05 | 0.037488 |  | 0.055758 | -0.0008721 | -5.7 | 0.90633 | 6.6323 | [5 empty] | -0.033601
-5.6 | -0.000835 | 0.84532 | -0.033601 | 6.5988 | 0.84532 | 0.79927 |  | 0.87502 | 0.90472 | 1e-05 | 0.072837 |  | 0.057658 | -0.000835 | -5.6 | 0.90472 | 6.5988 | [5 empty] | -0.033601
-5.5 | -0.000835 | 0.84196 | -0.033601 | 6.5653 | 0.84196 | 0.79504 |  | 0.87294 | 0.9031 | 1e-05 | 0.071224 |  | 0.059622 | -0.000835 | -5.5 | 0.9031 | 6.5653 | [5 empty] | -0.033601
-5.4 | -0.000835 | 0.8386 | 0 | 6.5318 | 0.8386 | 0.79081 |  | 0.87086 | 0.90149 | 1e-05 | 0.06961 |  | 0.061653 | -0.000835 | -5.4 | 0.90149 | 6.5318 | [5 empty] | -0.033601
-5.3 | -0.000835 | 0.8386 | -0.067203 | 6.4983 | 0.8386 | 0.79081 |  | 0.86878 | 0.89987 | 1e-05 | 0.067994 |  | 0.063753 | -0.000835 | -5.3 | 0.89987 | 6.4983 | [5 empty] | -0.033601
-5.2 | -0.000835 | 0.83188 | -0.033601 | 6.4648 | 0.83188 | 0.78237 |  | 0.8667 | 0.89826 | 1e-05 | 0.066378 |  | 0.065924 | -0.000835 | -5.2 | 0.89826 | 6.4648 | [5 empty] | -0.033601
-5.1 | -0.000835 | 0.82852 | -0.033601 | 6.4313 | 0.82852 | 0.77816 |  | 0.86462 | 0.89664 | 1e-05 | 0.064761 |  | 0.06817 | -0.000835 | -5.1 | 0.89664 | 6.4313 | [5 empty] | -0.036672
-5 | -0.000835 | 0.82516 | -0.033601 | 6.3978 | 0.82516 | 0.77396 |  | 0.86254 | 0.89502 | 1e-05 | 0.063143 |  | 0.070492 | -0.000835 | -5 | 0.89502 | 6.3978 | [5 empty] | -0.039744
-4.9 | -0.000835 | 0.8218 | -0.033601 | 6.3643 | 0.8218 | 0.76976 |  | 0.86046 | 0.8934 | 1e-05 | 0.061524 |  | 0.072893 | -0.000835 | -4.9 | 0.8934 | 6.3643 | [5 empty] | -0.042815
-4.8 | -0.000835 | 0.81844 | -0.033601 | 6.3308 | 0.81844 | 0.76557 |  | 0.85838 | 0.89178 | 1e-05 | 0.059904 |  | 0.075376 | -0.000835 | -4.8 | 0.89178 | 6.3308 | [5 empty] | -0.04894
-4.7 | -0.0007979 | 0.81508 | -0.033601 | 6.2973 | 0.81508 | 0.76138 |  | 0.8563 | 0.89016 | 1e-05 | 0.095244 |  | 0.077944 | -0.0007979 | -4.7 | 0.89016 | 6.2973 | [5 empty] | -0.045902
-4.6 | -0.000835 | 0.81172 | -0.067384 | 6.2638 | 0.81172 | 0.7572 |  | 0.85422 | 0.88854 | 1e-05 | 0.05666 |  | 0.080599 | -0.000835 | -4.6 | 0.88854 | 6.2638 | [5 empty] | -0.048973
-4.5 | -0.0007979 | 0.80498 | -0.067384 | 6.2303 | 0.80498 | 0.74883 |  | 0.85214 | 0.88692 | 1e-05 | 0.091999 |  | 0.083345 | -0.0007979 | -4.5 | 0.88692 | 6.2303 | [5 empty] | -0.04899
-4.4 | -0.0007979 | 0.79824 | -0.067384 | 6.1968 | 0.79824 | 0.74048 |  | 0.85006 | 0.88529 | 1e-05 | 0.090375 |  | 0.086184 | -0.0007979 | -4.4 | 0.88529 | 6.1968 | [5 empty] | -0.049006
-4.3 | -0.0007979 | 0.7915 | -0.067384 | 6.1633 | 0.7915 | 0.73216 |  | 0.84798 | 0.88367 | 1e-05 | 0.08875 |  | 0.089119 | -0.0007979 | -4.3 | 0.88367 | 6.1633 | [5 empty] | -0.049023
-4.2 | -0.0007979 | 0.78477 | -0.033782 | 6.1298 | 0.78477 | 0.72386 |  | 0.8459 | 0.88204 | 1e-05 | 0.087123 |  | 0.092155 | -0.0007979 | -4.2 | 0.88204 | 6.1298 | [5 empty] | -0.052094
-4.1 | -0.0007979 | 0.78139 | -0.067384 | 6.0963 | 0.78139 | 0.71971 |  | 0.84382 | 0.88042 | 1e-05 | 0.085496 |  | 0.095294 | -0.0007979 | -4.1 | 0.88042 | 6.0963 | [5 empty] | -0.055165
-4 | -0.0007606 | 0.77465 | -0.033782 | 6.0629 | 0.77465 | 0.71144 |  | 0.84174 | 0.87879 | 1e-05 | 0.12103 |  | 0.09854 | -0.0007606 | -4 | 0.87879 | 6.0629 | [5 empty] | -0.052094
-3.9 | -0.0007606 | 0.77127 | -0.033782 | 6.0294 | 0.77127 | 0.70731 |  | 0.83966 | 0.87716 | 1e-05 | 0.1194 |  | 0.1019 | -0.0007606 | -3.9 | 0.87716 | 6.0294 | [5 empty] | -0.049023
-3.8 | -0.0007606 | 0.76789 | -0.033782 | 5.9959 | 0.76789 | 0.70318 |  | 0.83758 | 0.87553 | 1e-05 | 0.11777 |  | 0.10537 | -0.0007606 | -3.8 | 0.87553 | 5.9959 | [5 empty] | -0.045952
-3.7 | -0.0007606 | 0.76451 | -0.067384 | 5.9624 | 0.76451 | 0.69906 |  | 0.8355 | 0.8739 | 1e-05 | 0.11614 |  | 0.10896 | -0.0007606 | -3.7 | 0.8739 | 5.9624 | [5 empty] | -0.042881
-3.6 | -0.0007606 | 0.75778 | -0.067384 | 5.9289 | 0.75778 | 0.69086 |  | 0.83342 | 0.87226 | 1e-05 | 0.11451 |  | 0.11267 | -0.0007606 | -3.6 | 0.87226 | 5.9289 | [5 empty] | -0.042864
-3.5 | -0.0007606 | 0.75104 | -0.033601 | 5.8954 | 0.75104 | 0.68268 |  | 0.83134 | 0.87063 | 1e-05 | 0.11287 |  | 0.11651 | -0.0007606 | -3.5 | 0.87063 | 5.8954 | [5 empty] | -0.039793
-3.4 | -0.0007606 | 0.74768 | -0.033601 | 5.8619 | 0.74768 | 0.67861 |  | 0.82926 | 0.869 | 1e-05 | 0.11124 |  | 0.12048 | -0.0007606 | -3.4 | 0.869 | 5.8619 | [5 empty] | -0.042839
-3.3 | -0.0007606 | 0.74432 | -0.033601 | 5.8284 | 0.74432 | 0.67455 |  | 0.82718 | 0.86736 | 1e-05 | 0.1096 |  | 0.12458 | -0.0007606 | -3.3 | 0.86736 | 5.8284 | [5 empty] | -0.045886
-3.2 | -0.0007606 | 0.74096 | -0.033601 | 5.7949 | 0.74096 | 0.67049 |  | 0.8251 | 0.86573 | 1e-05 | 0.10797 |  | 0.12882 | -0.0007606 | -3.2 | 0.86573 | 5.7949 | [5 empty] | -0.048932
-3.1 | -0.0007235 | 0.7376 | -0.033601 | 5.7614 | 0.7376 | 0.66644 |  | 0.82302 | 0.86409 | 1e-05 | 0.14329 |  | 0.13321 | -0.0007235 | -3.1 | 0.86409 | 5.7614 | [5 empty] | -0.045869
-3 | -0.0007235 | 0.73424 | -0.033601 | 5.7279 | 0.73424 | 0.66239 |  | 0.82094 | 0.86245 | 1e-05 | 0.14165 |  | 0.13775 | -0.0007235 | -3 | 0.86245 | 5.7279 | [5 empty] | -0.042806
-2.9 | -0.0007235 | 0.73088 | -0.067293 | 5.6944 | 0.73088 | 0.65835 |  | 0.81886 | 0.86081 | 1e-05 | 0.14001 |  | 0.14244 | -0.0007235 | -2.9 | 0.86081 | 5.6944 | [5 empty] | -0.045878
-2.8 | -0.0007235 | 0.72415 | -0.067293 | 5.6609 | 0.72415 | 0.65028 |  | 0.81678 | 0.85917 | 1e-05 | 0.13837 |  | 0.14729 | -0.0007235 | -2.8 | 0.85917 | 5.6609 | [5 empty] | -0.048949
-2.7 | -0.0007235 | 0.71742 | -0.067293 | 5.6274 | 0.71742 | 0.64224 |  | 0.8147 | 0.85753 | 1e-05 | 0.13673 |  | 0.15231 | -0.0007235 | -2.7 | 0.85753 | 5.6274 | [5 empty] | -0.05202
-2.6 | -0.0007235 | 0.71069 | -0.033692 | 5.5939 | 0.71069 | 0.63422 |  | 0.81262 | 0.85589 | 1e-05 | 0.13509 |  | 0.1575 | -0.0007235 | -2.6 | 0.85589 | 5.5939 | [5 empty] | -0.055091
-2.5 | -0.0007235 | 0.70732 | -0.033692 | 5.5604 | 0.70732 | 0.63021 |  | 0.81054 | 0.85424 | 1e-05 | 0.13345 |  | 0.16286 | -0.0007235 | -2.5 | 0.85424 | 5.5604 | [5 empty] | -0.061217
-2.4 | -0.0007235 | 0.70395 | -0.067384 | 5.5269 | 0.70395 | 0.62621 |  | 0.80846 | 0.8526 | 1e-05 | 0.1318 |  | 0.16841 | -0.0007235 | -2.4 | 0.8526 | 5.5269 | [5 empty] | -0.067343
-2.3 | -0.0006863 | 0.69721 | -0.067384 | 5.4934 | 0.69721 | 0.61824 |  | 0.80638 | 0.85095 | 1e-05 | 0.16722 |  | 0.17415 | -0.0006863 | -2.3 | 0.85095 | 5.4934 | [5 empty] | -0.067343
-2.2 | -0.0006863 | 0.69047 | -0.067384 | 5.4599 | 0.69047 | 0.61028 |  | 0.8043 | 0.84931 | 1e-05 | 0.16557 |  | 0.18008 | -0.0006863 | -2.2 | 0.84931 | 5.4599 | [5 empty] | -0.070405
-2.1 | -0.0006863 | 0.68374 | -0.067384 | 5.4264 | 0.68374 | 0.60235 |  | 0.80222 | 0.84766 | 1e-05 | 0.16392 |  | 0.18621 | -0.0006863 | -2.1 | 0.84766 | 5.4264 | [5 empty] | -0.073468
-2 | -0.0006863 | 0.677 | -0.10098 | 5.3929 | 0.677 | 0.59445 |  | 0.80014 | 0.84601 | 1e-05 | 0.16227 |  | 0.19256 | -0.0006863 | -2 | 0.84601 | 5.3929 | [5 empty] | -0.079586
-1.9 | -0.0006863 | 0.6669 | -0.10098 | 5.3594 | 0.6669 | 0.58266 |  | 0.79806 | 0.84436 | 1e-05 | 0.16062 |  | 0.19912 | -0.0006863 | -1.9 | 0.84436 | 5.3594 | [5 empty] | -0.088766
-1.8 | -0.0006491 | 0.6568 | -0.067293 | 5.3259 | 0.6568 | 0.57092 |  | 0.79598 | 0.84271 | 1e-05 | 0.19603 |  | 0.2059 | -0.0006491 | -1.8 | 0.84271 | 5.3259 | [5 empty] | -0.088758
-1.7 | -0.0006491 | 0.65007 | -0.10098 | 5.2924 | 0.65007 | 0.56314 |  | 0.7939 | 0.84106 | 1e-05 | 0.19438 |  | 0.21291 | -0.0006491 | -1.7 | 0.84106 | 5.2924 | [5 empty] | -0.091813
-1.6 | -0.0006491 | 0.63997 | -0.10098 | 5.2589 | 0.63997 | 0.5515 |  | 0.79182 | 0.8394 | 1e-05 | 0.19273 |  | 0.22017 | -0.0006491 | -1.6 | 0.8394 | 5.2589 | [5 empty] | -0.094867
-1.5 | -0.0006491 | 0.62987 | -0.10098 | 5.2254 | 0.62987 | 0.53993 |  | 0.78974 | 0.83775 | 1e-05 | 0.19108 |  | 0.22767 | -0.0006491 | -1.5 | 0.83775 | 5.2254 | [5 empty] | -0.10098
-1.4 | -0.000612 | 0.61978 | -0.13468 | 5.1919 | 0.61978 | 0.52842 |  | 0.78766 | 0.83609 | 1e-05 | 0.22638 |  | 0.23542 | -0.000612 | -1.4 | 0.83609 | 5.1919 | [5 empty] | -0.09793
-1.3 | -0.000612 | 0.60631 | -0.067293 | 5.1584 | 0.60631 | 0.51317 |  | 0.78558 | 0.83444 | 1e-05 | 0.22472 |  | 0.24344 | -0.000612 | -1.3 | 0.83444 | 5.1584 | [5 empty] | -0.097939
-1.2 | -0.000612 | 0.59958 | -0.10098 | 5.125 | 0.59958 | 0.50559 |  | 0.7835 | 0.83278 | 1e-05 | 0.22307 |  | 0.25173 | -0.000612 | -1.2 | 0.83278 | 5.125 | [5 empty] | -0.10406
-1.1 | -0.0005748 | 0.58948 | -0.10098 | 5.0915 | 0.58948 | 0.49426 |  | 0.78142 | 0.83112 | 1e-05 | 0.25847 |  | 0.26031 | -0.0005748 | -1.1 | 0.83112 | 5.0915 | [5 empty] | -0.10406
-1 | -0.0005748 | 0.57938 | -0.13468 | 5.058 | 0.57938 | 0.48301 |  | 0.77934 | 0.82946 | 1e-05 | 0.25681 |  | 0.26917 | -0.0005748 | -1 | 0.82946 | 5.058 | [5 empty] | -0.10406
-0.9 | -0.0005748 | 0.56591 | -0.067384 | 5.0245 | 0.56591 | 0.4681 |  | 0.77726 | 0.8278 | 1e-05 | 0.25515 |  | 0.27834 | -0.0005748 | -0.9 | 0.8278 | 5.0245 | [5 empty] | -0.10714
-0.8 | -0.0005376 | 0.55918 | -0.10108 | 4.991 | 0.55918 | 0.46068 |  | 0.77518 | 0.82614 | 1e-05 | 0.29055 |  | 0.28782 | -0.0005376 | -0.8 | 0.82614 | 4.991 | [5 empty] | -0.10101
-0.7 | -0.0005748 | 0.54907 | -0.13468 | 4.9575 | 0.54907 | 0.44961 |  | 0.7731 | 0.82448 | 1e-05 | 0.25182 |  | 0.29763 | -0.0005748 | -0.7 | 0.82448 | 4.9575 | [5 empty] | -0.1102
-0.6 | -0.0005376 | 0.5356 | -0.10098 | 4.924 | 0.5356 | 0.43497 |  | 0.77102 | 0.82281 | 1e-05 | 0.28722 |  | 0.30777 | -0.0005376 | -0.6 | 0.82281 | 4.924 | [5 empty] | -0.11326
-0.5 | -0.0005376 | 0.5255 | -0.10098 | 4.8905 | 0.5255 | 0.42407 |  | 0.76894 | 0.82115 | 1e-05 | 0.28556 |  | 0.31825 | -0.0005376 | -0.5 | 0.82115 | 4.8905 | [5 empty] | -0.11632
-0.4 | -0.0005004 | 0.5154 | -0.13477 | 4.857 | 0.5154 | 0.41323 |  | 0.76686 | 0.81948 | 1e-05 | 0.32095 |  | 0.32909 | -0.0005004 | -0.4 | 0.81948 | 4.857 | [5 empty] | -0.11632
-0.3 | -0.0005376 | 0.50193 | -0.067293 | 4.8235 | 0.50193 | 0.39889 |  | 0.76478 | 0.81781 | 1e-05 | 0.28222 |  | 0.3403 | -0.0005376 | -0.3 | 0.81781 | 4.8235 | [5 empty] | -0.12856
-0.2 | -0.0005004 | 0.4952 | -0.16837 | 4.79 | 0.4952 | 0.39178 |  | 0.7627 | 0.81615 | 1e-05 | 0.31761 |  | 0.35189 | -0.0005004 | -0.2 | 0.81615 | 4.79 | [5 empty] | -0.13468
-0.1 | -0.0004633 | 0.47836 | -0.13468 | 4.7565 | 0.47836 | 0.37412 |  | 0.76062 | 0.81448 | 1e-05 | 0.35291 |  | 0.36388 | -0.0004633 | -0.1 | 0.81448 | 4.7565 | [5 empty] | -0.13774
0 | -0.0004633 | 0.46489 | -0.13468 | 4.723 | 0.46489 | 0.36014 |  | 0.75855 | 0.8128 | 1e-05 | 0.35124 |  | 0.37628 | -0.0004633 | 0 | 0.8128 | 4.723 | [5 empty] | -0.14387
0.1 | -0.0004633 | 0.45143 | -0.13459 | 4.6895 | 0.45143 | 0.3463 |  | 0.75647 | 0.81113 | 1e-05 | 0.34956 |  | 0.38909 | -0.0004633 | 0.1 | 0.81113 | 4.6895 | [5 empty] | -0.14999
0.2 | -0.000426 | 0.43797 | -0.20206 | 4.656 | 0.43797 | 0.3326 |  | 0.75439 | 0.80946 | 1e-05 | 0.38505 |  | 0.40235 | -0.000426 | 0.2 | 0.80946 | 4.656 | [5 empty] | -0.1561
0.3 | -0.0004633 | 0.41776 | -0.16837 | 4.6225 | 0.41776 | 0.3123 |  | 0.75231 | 0.80778 | 1e-05 | 0.34622 |  | 0.41605 | -0.0004633 | 0.3 | 0.80778 | 4.6225 | [5 empty] | -0.16834
0.4 | -0.0003889 | 0.40092 | -0.16837 | 4.589 | 0.40092 | 0.29563 |  | 0.75023 | 0.80611 | 1e-05 | 0.41866 |  | 0.43023 | -0.0003889 | 0.4 | 0.80611 | 4.589 | [5 empty] | -0.16834
0.5 | -0.0003889 | 0.38409 | -0.16837 | 4.5555 | 0.38409 | 0.27919 |  | 0.74815 | 0.80443 | 1e-05 | 0.41699 |  | 0.44488 | -0.0003889 | 0.5 | 0.80443 | 4.5555 | [5 empty] | -0.17446
0.6 | -0.0003889 | 0.36725 | -0.16837 | 4.522 | 0.36725 | 0.263 |  | 0.74607 | 0.80275 | 1e-05 | 0.41531 |  | 0.46003 | -0.0003889 | 0.6 | 0.80275 | 4.522 | [5 empty] | -0.18058
0.7 | -0.0003518 | 0.35041 | -0.20197 | 4.4885 | 0.35041 | 0.24704 |  | 0.74399 | 0.80108 | 1e-05 | 0.45059 |  | 0.47571 | -0.0003518 | 0.7 | 0.80108 | 4.4885 | [5 empty] | -0.18671
0.8 | -0.0003145 | 0.33022 | -0.20197 | 4.455 | 0.33022 | 0.22824 |  | 0.74191 | 0.7994 | 1e-05 | 0.48607 |  | 0.49191 | -0.0003145 | 0.8 | 0.7994 | 4.455 | [5 empty] | -0.1867
0.9 | -0.0003145 | 0.31002 | -0.16837 | 4.4215 | 0.31002 | 0.20982 |  | 0.73983 | 0.79771 | 1e-05 | 0.48439 |  | 0.50867 | -0.0003145 | 0.9 | 0.79771 | 4.4215 | [5 empty] | -0.18976
1 | -0.0002774 | 0.29318 | -0.20197 | 4.388 | 0.29318 | 0.19477 |  | 0.73775 | 0.79603 | 1e-05 | 0.51967 |  | 0.52599 | -0.0002774 | 1 | 0.79603 | 4.388 | [5 empty] | -0.19282
1.1 | -0.0002774 | 0.27299 | -0.20197 | 4.3545 | 0.27299 | 0.17709 |  | 0.73567 | 0.79435 | 1e-05 | 0.51798 |  | 0.54391 | -0.0002774 | 1.1 | 0.79435 | 4.3545 | [5 empty] | -0.19894
1.2 | -0.0002774 | 0.25279 | -0.20206 | 4.321 | 0.25279 | 0.15984 |  | 0.73359 | 0.79266 | 1e-05 | 0.5163 |  | 0.56244 | -0.0002774 | 1.2 | 0.79266 | 4.321 | [5 empty] | -0.20811
1.3 | -0.000203 | 0.23258 | -0.20197 | 4.2875 | 0.23258 | 0.14303 |  | 0.73151 | 0.79098 | 1e-05 | 0.58874 |  | 0.5816 | -0.000203 | 1.3 | 0.79098 | 4.2875 | [5 empty] | -0.21117
1.4 | -0.0002403 | 0.21239 | -0.20197 | 4.254 | 0.21239 | 0.12672 |  | 0.72943 | 0.78929 | 1e-05 | 0.54989 |  | 0.60141 | -0.0002403 | 1.4 | 0.78929 | 4.254 | [5 empty] | -0.22035
1.5 | -0.000203 | 0.19219 | -0.20206 | 4.2205 | 0.19219 | 0.11091 |  | 0.72735 | 0.7876 | 1e-05 | 0.58536 | -0.53553 | 0.62189 | -0.000203 | 1.5 | 0.7876 | 4.2205 | [5 empty] | -0.22951
1.6 | -0.0001659 | 0.17198 | -0.23566 | 4.1871 | 0.17198 | 0.095641 |  | 0.72527 | 0.78591 | 1e-05 | 0.62063 | -0.47702 | 0.64308 | -0.0001659 | 1.6 | 0.78591 | 4.1871 | [5 empty] | -0.23564
1.7 | -0.0001659 | 0.14842 | -0.26926 | 4.1536 | 0.14842 | 0.078579 |  | 0.72319 | 0.78422 | 1e-05 | 0.61894 | -0.47975 | 0.66498 | -0.0001659 | 1.7 | 0.78422 | 4.1536 | [5 empty] | -0.24176
1.8 | -0.0001287 | 0.12149 | -0.23557 | 4.1201 | 0.12149 | 0.060171 |  | 0.72111 | 0.78253 | 1e-05 | 0.65431 | -0.42418 | 0.68764 | -0.0001287 | 1.8 | 0.78253 | 4.1201 | [5 empty] | -0.25093
1.9 | -9.15e-05 | 0.097933 | -0.30295 | 4.0866 | 0.097933 | 0.045141 |  | 0.71903 | 0.78083 | 1e-05 | 0.68968 | -0.37153 | 0.71106 | -9.15e-05 | 1.9 | 0.78083 | 4.0866 | [5 empty] | -0.25704
2 | -9.15e-05 | 0.067637 | -0.26917 | 4.0531 | 0.067637 | 0.027558 |  | 0.71695 | 0.77914 | 1e-05 | 0.68798 | -0.37399 | 0.73528 | -9.15e-05 | 2 | 0.77914 | 4.0531 | [5 empty] | -0.26927
2.1 | -5.43e-05 | 0.04072 | -0.26935 | 4.0196 | 0.04072 | 0.014009 |  | 0.71487 | 0.77744 | 1e-05 | 0.72335 | -0.32387 | 0.76033 | -5.43e-05 | 2.1 | 0.77744 | 4.0196 | [5 empty] | -0.2815
2.2 | -1.72e-05 | 0.013785 | -0.26935 | 3.9861 | 0.013785 | 0.0033052 |  | 0.71279 | 0.77575 | 1e-05 | 0.75861 | -0.27627 | 0.78623 | -1.72e-05 | 2.2 | 0.77575 | 3.9861 | [5 empty] | -0.28762
2.3 | 1.99e-05 | -0.013151 | -0.30286 | 3.9526 |  |  |  | 0.71071 | 0.77405 | 0.019826 | 0.79387 | -0.23083 | 0.81301 | 1.99e-05 | 2.3 | 0.77405 | 3.9526 | [5 empty] | -0.2968
2.4 | 5.71e-05 | -0.043437 | -0.26917 | 3.9191 |  |  |  | 0.70863 | 0.77235 | 0.056887 | 0.82924 | -0.18725 | 0.8407 | 5.71e-05 | 2.4 | 0.77235 | 3.9191 | [5 empty] | -0.30599
2.5 | 9.42e-05 | -0.070354 | -0.33656 | 3.8856 |  |  |  | 0.70655 | 0.77065 | 0.093848 | 0.8645 | -0.14561 | 0.86934 | 9.42e-05 | 2.5 | 0.77065 | 3.8856 | [5 empty] | -0.31213
2.6 | 9.42e-05 | -0.10401 | -0.33656 | 3.8521 |  |  |  | 0.70447 | 0.76895 | 0.093848 | 0.86279 | -0.14758 | 0.89895 | 9.42e-05 | 2.6 | 0.76895 | 3.8521 | [5 empty] | -0.32439
2.7 | 0.0001315 | -0.13767 | -0.30295 | 3.8186 |  |  |  | 0.70239 | 0.76724 | 0.13101 | 0.89825 | -0.10731 | 0.92957 | 0.0001315 | 2.7 | 0.76724 | 3.8186 | [5 empty] | -0.33663
2.8 | 0.0001315 | -0.16796 | -0.37034 | 3.7851 |  |  |  | 0.70031 | 0.76554 | 0.13101 | 0.89655 | -0.1092 | 0.96124 | 0.0001315 | 2.8 | 0.76554 | 3.7851 | [5 empty] | -0.35498
2.9 | 0.0002057 | -0.20499 | -0.33665 | 3.7516 |  |  |  | 0.69823 | 0.76383 | 0.20493 | 0.96876 | -0.031734 | 0.99398 | 0.0002057 | 2.9 | 0.76383 | 3.7516 | [5 empty] | -0.37031
3 | 0.0002057 | -0.23866 | -0.37043 | 3.7181 |  |  |  | 0.69615 | 0.76213 | 0.20493 | 0.96706 | -0.033498 | 1.0278 | 0.0002057 | 3 | 0.76213 | 3.7181 | [5 empty] | -0.39175
3.1 | 0.0002801 | -0.2757 | -0.40403 | 3.6846 |  |  |  | 0.69407 | 0.76042 | 0.27905 | 1.0395 | 0.038711 | 1.0629 | 0.0002801 | 3.1 | 0.76042 | 3.6846 | [5 empty] | -0.40401
3.2 | 0.0003173 | -0.31611 | -0.40403 | 3.6511 |  |  |  | 0.69199 | 0.75871 | 0.31611 | 1.0748 | 0.072155 | 1.0991 | 0.0003173 | 3.2 | 0.75871 | 3.6511 | [5 empty] | -0.41932
3.3 | 0.0003173 | -0.35651 | -0.47123 | 3.6176 |  |  |  | 0.68991 | 0.757 | 0.31611 | 1.0731 | 0.070562 | 1.1365 | 0.0003173 | 3.3 | 0.757 | 3.6176 | [5 empty] | -0.44075
3.4 | 0.0004288 | -0.40363 | -0.47141 | 3.5841 |  |  |  | 0.68783 | 0.75528 | 0.4272 | 1.1825 | 0.16762 | 1.1752 | 0.0004288 | 3.4 | 0.75528 | 3.5841 | [5 empty] | -0.453
3.5 | 0.0004288 | -0.45077 | -0.50502 | 3.5506 |  |  |  | 0.68575 | 0.75357 | 0.4272 | 1.1808 | 0.16617 | 1.2152 | 0.0004288 | 3.5 | 0.75357 | 3.5506 | [5 empty] | -0.47136
3.6 | 0.0005032 | -0.50127 | -0.47141 | 3.5171 |  |  |  | 0.68367 | 0.75186 | 0.50132 | 1.2532 | 0.22568 | 1.2566 | 0.0005032 | 3.6 | 0.75186 | 3.5171 | [5 empty] | -0.48665
3.7 | 0.0005403 | -0.54842 | -0.50502 | 3.4836 |  |  |  | 0.68159 | 0.75014 | 0.53828 | 1.2884 | 0.25342 | 1.2994 | 0.0005403 | 3.7 | 0.75014 | 3.4836 | [5 empty] | -0.50196
3.8 | 0.0005776 | -0.59892 | -0.53871 | 3.4501 |  |  |  | 0.67951 | 0.74842 | 0.57544 | 1.3239 | 0.28055 | 1.3437 | 0.0005776 | 3.8 | 0.74842 | 3.4501 | [5 empty] | -0.51725
3.9 | 0.0006518 | -0.65279 | -0.50502 | 3.4166 |  |  |  | 0.67743 | 0.7467 | 0.64936 | 1.3961 | 0.33366 | 1.3895 | 0.0006518 | 3.9 | 0.7467 | 3.4166 | [5 empty] | -0.52951
4 | 0.0007262 | -0.70329 | -0.53871 | 3.3831 |  |  |  | 0.67535 | 0.74498 | 0.72349 | 1.4685 | 0.38422 | 1.4368 | 0.0007262 | 4 | 0.74498 | 3.3831 | [5 empty] | -0.54175
4.1 | 0.0007633 | -0.75716 | -0.53862 | 3.3496 |  |  |  | 0.67327 | 0.74326 | 0.76045 | 1.5037 | 0.40794 | 1.4858 | 0.0007633 | 4.1 | 0.74326 | 3.3496 | [5 empty] | -0.554
4.2 | 0.0008006 | -0.81102 | -0.5724 | 3.3161 |  |  |  | 0.67119 | 0.74154 | 0.79761 | 1.5391 | 0.43123 | 1.5364 | 0.0008006 | 4.2 | 0.74154 | 3.3161 | [5 empty] | -0.57235
4.3 | 0.0008749 | -0.86826 | -0.57222 | 3.2826 |  |  |  | 0.66911 | 0.73981 | 0.87163 | 1.6114 | 0.47713 | 1.5887 | 0.0008749 | 4.3 | 0.73981 | 3.2826 | [5 empty] | -0.59071
4.4 | 0.0009121 | -0.92548 | -0.60609 | 3.2492 |  |  |  | 0.66703 | 0.73809 | 0.90869 | 1.6468 | 0.49882 | 1.6428 | 0.0009121 | 4.4 | 0.73809 | 3.2492 | [5 empty] | -0.60907
4.5 | 0.0009864 | -0.98609 | -0.606 | 3.2157 |  |  |  | 0.66495 | 0.73636 | 0.98271 | 1.7191 | 0.54179 | 1.6988 | 0.0009864 | 4.5 | 0.73636 | 3.2157 | [5 empty] | -0.62743
4.6 | 0.0010236 | -1.0467 | -0.63978 | 3.1822 |  |  |  | 0.66287 | 0.73463 | 1.0198 | 1.7544 | 0.56213 | 1.7566 | 0.0010236 | 4.6 | 0.73463 | 3.1822 | [5 empty] | -0.65192
4.7 | 0.0010979 | -1.1107 | -0.67329 | 3.1487 |  |  |  | 0.66079 | 0.7329 | 1.0938 | 1.8267 | 0.60251 | 1.8165 | 0.0010979 | 4.7 | 0.7329 | 3.1487 | [5 empty] | -0.6703
4.8 | 0.0011723 | -1.178 | -0.70699 | 3.1152 |  |  |  | 0.65871 | 0.73117 | 1.1679 | 1.8991 | 0.64138 | 1.8784 | 0.0011723 | 4.8 | 0.73117 | 3.1152 | [5 empty] | -0.69172
4.9 | 0.0012094 | -1.2487 | -0.74068 | 3.0817 |  |  |  | 0.65663 | 0.72944 | 1.2049 | 1.9343 | 0.65976 | 1.9423 | 0.0012094 | 4.9 | 0.72944 | 3.0817 | [5 empty] | -0.72234
5 | 0.001321 | -1.3228 | -0.70699 | 3.0482 |  |  |  | 0.65455 | 0.72771 | 1.3161 | 2.0438 | 0.7148 | 2.0085 | 0.001321 | 5 | 0.72771 | 3.0482 | [5 empty] | -0.74376
5.1 | 0.0013953 | -1.3935 | -0.80806 | 3.0147 |  |  |  | 0.65247 | 0.72597 | 1.3901 | 2.1161 | 0.74956 | 2.0769 | 0.0013953 | 5.1 | 0.72597 | 3.0147 | [5 empty] | -0.77131
5.2 | 0.0014697 | -1.4743 | -0.74077 | 2.9812 |  |  |  | 0.65039 | 0.72424 | 1.4642 | 2.1884 | 0.78319 | 2.1477 | 0.0014697 | 5.2 | 0.72424 | 2.9812 | [5 empty] | -0.79886
5.3 | 0.001544 | -1.5483 | -0.80806 | 2.9477 |  |  |  | 0.64831 | 0.7225 | 1.5382 | 2.2607 | 0.81569 | 2.2208 | 0.001544 | 5.3 | 0.7225 | 2.9477 | [5 empty] | -0.82335
5.4 | 0.0016555 | -1.6292 | -0.90905 | 2.9142 |  |  |  | 0.64623 | 0.72076 | 1.6493 | 2.3701 | 0.86292 | 2.2965 | 0.0016555 | 5.4 | 0.72076 | 2.9142 | [5 empty] | -0.8509
5.5 | 0.0017299 | -1.7201 | -0.84175 | 2.8807 |  |  |  | 0.64415 | 0.71902 | 1.7234 | 2.4425 | 0.893 | 2.3747 | 0.0017299 | 5.5 | 0.71902 | 2.8807 | [5 empty] | -0.87233
5.6 | 0.001767 | -1.8042 | -0.90905 | 2.8472 |  |  |  | 0.64207 | 0.71728 | 1.7604 | 2.4777 | 0.90732 | 2.4556 | 0.001767 | 5.6 | 0.71728 | 2.8472 | [5 empty] | -0.90601
5.7 | 0.0019158 | -1.8951 | -0.94274 | 2.8137 |  |  |  | 0.63999 | 0.71553 | 1.9086 | 2.6242 | 0.96477 | 2.5392 | 0.0019158 | 5.7 | 0.71553 | 2.8137 | [5 empty] | -0.92743
5.8 | 0.0019158 | -1.9894 | -0.94274 | 2.7802 |  |  |  | 0.63791 | 0.71379 | 1.9086 | 2.6224 | 0.9641 | 2.6257 | 0.0019158 | 5.8 | 0.71379 | 2.7802 | [5 empty] | -0.96721
5.9 | 0.0020645 | -2.0837 | -1.01 | 2.7467 |  |  |  | 0.63583 | 0.71204 | 2.0568 | 2.7688 | 1.0184 | 2.7152 | 0.0020645 | 5.9 | 0.71204 | 2.7467 | [5 empty] | -1.0009
6 | 0.0022131 | -2.1847 | -0.97643 | 2.7132 |  |  |  | 0.63375 | 0.7103 | 2.2048 | 2.9151 | 1.0699 | 2.8077 | 0.0022131 | 6 | 0.7103 | 2.7132 | [5 empty] | -1.0223
6.1 | 0.0022504 | -2.2823 | -1.0774 | 2.6797 |  |  |  | 0.63167 | 0.70855 | 2.242 | 2.9505 | 1.082 | 2.9033 | 0.0022504 | 6.1 | 0.70855 | 2.6797 | [5 empty] | -1.059
6.2 | 0.002399 | -2.3901 | -1.0437 | 2.6462 |  |  |  | 0.62959 | 0.7068 | 2.39 | 3.0968 | 1.1304 | 3.0022 | 0.002399 | 6.2 | 0.7068 | 2.6462 | [5 empty] | -1.0927
6.3 | 0.0025106 | -2.4944 | -1.1783 | 2.6127 |  |  |  | 0.62751 | 0.70504 | 2.5012 | 3.2063 | 1.1651 | 3.1045 | 0.0025106 | 6.3 | 0.70504 | 2.6127 | [5 empty] | -1.1264
6.4 | 0.0025849 | -2.6123 | -1.1784 | 2.5792 |  |  |  | 0.62543 | 0.70329 | 2.5752 | 3.2785 | 1.1874 | 3.2102 | 0.0025849 | 6.4 | 0.70329 | 2.5792 | [5 empty] | -1.1723
6.5 | 0.0027707 | -2.7301 | -1.1448 | 2.5457 |  |  |  | 0.62335 | 0.70154 | 2.7603 | 3.4619 | 1.2418 | 3.3196 | 0.0027707 | 6.5 | 0.70154 | 2.5457 | [5 empty] | -1.209
6.6 | 0.002808 | -2.8446 | -1.2457 | 2.5122 |  |  |  | 0.62127 | 0.69978 | 2.7975 | 3.4973 | 1.252 | 3.4326 | 0.002808 | 6.6 | 0.69978 | 2.5122 | [5 empty] | -1.2641
6.7 | 0.0029566 | -2.9692 | -1.2795 | 2.4787 |  |  |  | 0.61919 | 0.69802 | 2.9456 | 3.6436 | 1.293 | 3.5496 | 0.0029566 | 6.7 | 0.69802 | 2.4787 | [5 empty] | -1.3131
6.8 | 0.0030682 | -3.0971 | -1.313 | 2.4452 |  |  |  | 0.61711 | 0.69626 | 3.0567 | 3.753 | 1.3226 | 3.6705 | 0.0030682 | 6.8 | 0.69626 | 2.4452 | [5 empty] | -1.3712
6.9 | 0.0032168 | -3.2284 | -1.4478 | 2.4117 |  |  |  | 0.61503 | 0.6945 | 3.2048 | 3.8993 | 1.3608 | 3.7955 | 0.0032168 | 6.9 | 0.6945 | 2.4117 | [5 empty] | -1.4141
7 | 0.0033656 | -3.3732 | -1.4142 | 2.3782 |  |  |  | 0.61295 | 0.69274 | 3.353 | 4.0458 | 1.3977 | 3.9248 | 0.0033656 | 7 | 0.69274 | 2.3782 | [5 empty] | -1.4569
7.1 | 0.0034771 | -3.5146 | -1.5823 | 2.3447 |  |  |  | 0.61087 | 0.69098 | 3.4641 | 4.1551 | 1.4243 | 4.0585 | 0.0034771 | 7.1 | 0.69098 | 2.3447 | [5 empty] | -1.5212
7.2 | 0.0036258 | -3.6728 | -1.6161 | 2.3113 |  |  |  | 0.60879 | 0.68921 | 3.6123 | 4.3015 | 1.459 | 4.1967 | 0.0036258 | 7.2 | 0.68921 | 2.3113 | [5 empty] | -1.5671
7.3 | 0.0038117 | -3.8344 | -1.6834 | 2.2778 |  |  |  | 0.60671 | 0.68744 | 3.7975 | 4.4849 | 1.5007 | 4.3397 | 0.0038117 | 7.3 | 0.68744 | 2.2778 | [5 empty] | -1.6222
7.4 | 0.0039603 | -4.0028 | -1.6498 | 2.2443 |  |  |  | 0.60463 | 0.68568 | 3.9455 | 4.6312 | 1.5328 | 4.4875 | 0.0039603 | 7.4 | 0.68568 | 2.2443 | [5 empty] | -1.6834
7.5 | 0.0041834 | -4.1678 | -1.6497 | 2.2108 |  |  |  | 0.60255 | 0.68391 | 4.1678 | 4.8517 | 1.5793 | 4.6404 | 0.0041834 | 7.5 | 0.68391 | 2.2108 | [5 empty] | -1.7324
7.6 | 0.0043321 | -4.3327 | -1.8518 | 2.1773 |  |  |  | 0.60047 | 0.68213 | 4.3159 | 4.9981 | 1.609 | 4.7985 | 0.0043321 | 7.6 | 0.68213 | 2.1773 | [5 empty] | -1.7967
7.7 | 0.0045551 | -4.5179 | -1.7507 | 2.1438 |  |  |  | 0.59839 | 0.68036 | 4.5381 | 5.2184 | 1.6522 | 4.9619 | 0.0045551 | 7.7 | 0.68036 | 2.1438 | [5 empty] | -1.8456
7.8 | 0.004741 | -4.693 | -1.8854 | 2.1103 |  |  |  | 0.59631 | 0.67859 | 4.7233 | 5.4019 | 1.6867 | 5.1309 | 0.004741 | 7.8 | 0.67859 | 2.1103 | [5 empty] | -1.8946
7.9 | 0.0049269 | -4.8815 | -1.9866 | 2.0768 |  |  |  | 0.59423 | 0.67681 | 4.9085 | 5.5853 | 1.7201 | 5.3057 | 0.0049269 | 7.9 | 0.67681 | 2.0768 | [5 empty] | -1.9436
8 | 0.0050384 | -5.0802 | -1.9865 | 2.0433 |  |  |  | 0.59215 | 0.67503 | 5.0196 | 5.6946 | 1.7395 | 5.4864 | 0.0050384 | 8 | 0.67503 | 2.0433 | [5 empty] | -2.0079
8.1 | 0.0051871 | -5.2788 | -2.121 | 2.0098 |  |  |  | 0.59007 | 0.67326 | 5.1677 | 5.841 | 1.7649 | 5.6733 | 0.0051871 | 8.1 | 0.67326 | 2.0098 | [5 empty] | -2.0813
8.2 | 0.0055217 | -5.4909 | -2.1211 | 1.9763 |  |  |  | 0.58799 | 0.67147 | 5.5011 | 6.1725 | 1.8201 | 5.8666 | 0.0055217 | 8.2 | 0.67147 | 1.9763 | [5 empty] | -2.118
8.3 | 0.0055588 | -5.703 | -2.1547 | 1.9428 |  |  |  | 0.58591 | 0.66969 | 5.538 | 6.2077 | 1.8258 | 6.0664 | 0.0055588 | 8.3 | 0.66969 | 1.9428 | [5 empty] | -2.2007
8.4 | 0.0058934 | -5.9185 | -2.2221 | 1.9093 |  |  |  | 0.58383 | 0.66791 | 5.8714 | 6.5393 | 1.8778 | 6.2731 | 0.0058934 | 8.4 | 0.66791 | 1.9093 | [5 empty] | -2.2588
8.5 | 0.0061537 | -6.1407 | -2.3568 | 1.8758 |  |  |  | 0.58175 | 0.66612 | 6.1307 | 6.7968 | 1.9165 | 6.4868 | 0.0061537 | 8.5 | 0.66612 | 1.8758 | [5 empty] | -2.3078
8.6 | 0.0063767 | -6.3764 | -2.4578 | 1.8423 |  |  |  | 0.57967 | 0.66434 | 6.3529 | 7.0172 | 1.9484 | 6.7077 | 0.0063767 | 8.6 | 0.66434 | 1.8423 | [5 empty] | -2.3782
8.7 | 0.006674 | -6.6222 | -2.2557 | 1.8088 |  |  |  | 0.57759 | 0.66255 | 6.6491 | 7.3116 | 1.9895 | 6.9362 | 0.006674 | 8.7 | 0.66255 | 1.8088 | [5 empty] | -2.4241
8.8 | 0.0068971 | -6.8478 | -2.6598 | 1.7753 |  |  |  | 0.57551 | 0.66076 | 6.8713 | 7.5321 | 2.0192 | 7.1725 | 0.0068971 | 8.8 | 0.66076 | 1.7753 | [5 empty] | -2.4884
8.9 | 0.0071201 | -7.1138 | -2.5251 | 1.7418 |  |  |  | 0.57343 | 0.65897 | 7.0935 | 7.7525 | 2.048 | 7.4168 | 0.0071201 | 8.9 | 0.65897 | 1.7418 | [5 empty] | -2.5618
9 | 0.0073804 | -7.3663 | -2.5251 | 1.7083 |  |  |  | 0.57135 | 0.65717 | 7.3528 | 8.01 | 2.0807 | 7.6695 | 0.0073804 | 9 | 0.65717 | 1.7083 | [5 empty] | -2.62
9.1 | 0.0076406 | -7.6188 | -2.7607 | 1.6748 |  |  |  | 0.56927 | 0.65538 | 7.6121 | 8.2674 | 2.1123 | 7.9307 | 0.0076406 | 9.1 | 0.65538 | 1.6748 | [5 empty] | -2.6934
9.2 | 0.0079008 | -7.8948 | -2.6262 | 1.6413 |  |  |  | 0.56719 | 0.65358 | 7.8713 | 8.5249 | 2.143 | 8.2009 | 0.0079008 | 9.2 | 0.65358 | 1.6413 | [5 empty] | -2.7577
9.3 | 0.0080123 | -8.1575 | -2.8281 | 1.6078 |  |  |  | 0.56511 | 0.65178 | 7.9824 | 8.6341 | 2.1557 | 8.4802 | 0.0080123 | 9.3 | 0.65178 | 1.6078 | [5 empty] | -2.8557
9.4 | 0.0084956 | -8.4403 | -2.9629 | 1.5743 |  |  |  | 0.56304 | 0.64998 | 8.4639 | 9.1138 | 2.2098 | 8.7691 | 0.0084956 | 9.4 | 0.64998 | 1.5743 | [5 empty] | -2.9291
9.5 | 0.0086814 | -8.7366 | -2.8617 | 1.5408 |  |  |  | 0.56096 | 0.64818 | 8.649 | 9.2971 | 2.2297 | 9.0678 | 0.0086814 | 9.5 | 0.64818 | 1.5408 | [5 empty] | -3.0118
9.6 | 0.0089417 | -9.0227 | -3.1648 | 1.5073 |  |  |  | 0.55888 | 0.64638 | 8.9083 | 9.5547 | 2.257 | 9.3767 | 0.0089417 | 9.6 | 0.64638 | 1.5073 | [5 empty] | -3.1128
9.7 | 0.0094249 | -9.3392 | -3.1649 | 1.4738 |  |  |  | 0.5568 | 0.64457 | 9.3897 | 10.0343 | 2.306 | 9.6961 | 0.0094249 | 9.7 | 0.64457 | 1.4738 | [5 empty] | -3.1954
9.8 | 0.0095736 | -9.6557 | -3.3331 | 1.4403 |  |  |  | 0.55472 | 0.64277 | 9.5378 | 10.1806 | 2.3205 | 10.0264 | 0.0095736 | 9.8 | 0.64277 | 1.4403 | [5 empty] | -3.3056
9.9 | 0.01002 | -9.989 | -3.4678 | 1.4068 |  |  |  | 0.55264 | 0.64096 | 9.9823 | 10.6232 | 2.363 | 10.3679 | 0.01002 | 9.9 | 0.64096 | 1.4068 | [5 empty] | -3.3944
10 | 0.010391 | -10.3358 | -3.4342 | 1.3734 |  |  |  | 0.55056 | 0.63915 | 10.3527 | 10.9918 | 2.3972 | 10.7211 | 0.010391 | 10 | 0.63915 | 1.3734 | [5 empty] | -3.4739
10.1 | 0.01054 | -10.6792 | -3.6362 | 1.3399 |  |  |  | 0.54848 | 0.63734 | 10.5007 | 11.1381 | 2.4104 | 11.0863 | 0.01054 | 10.1 | 0.63734 | 1.3399 | [5 empty] | -3.5841
10.2 | 0.011135 | -11.0428 | -3.6699 | 1.3064 |  |  |  | 0.5464 | 0.63552 | 11.0933 | 11.7288 | 2.462 | 11.4639 | 0.011135 | 10.2 | 0.63552 | 1.3064 | [5 empty] | -3.6637
10.3 | 0.011395 | -11.4098 | -3.8382 | 1.2729 |  |  |  | 0.54432 | 0.63371 | 11.3526 | 11.9863 | 2.4838 | 11.8544 | 0.011395 | 10.3 | 0.63371 | 1.2729 | [5 empty] | -3.7861
10.4 | 0.011692 | -11.7936 | -3.8046 | 1.2394 |  |  |  | 0.54224 | 0.63189 | 11.6488 | 12.2807 | 2.508 | 12.2582 | 0.011692 | 10.4 | 0.63189 | 1.2394 | [5 empty] | -3.8963
10.5 | 0.012325 | -12.1741 | -3.8382 | 1.2059 |  |  |  | 0.54016 | 0.63007 | 12.2785 | 12.9085 | 2.5579 | 12.6758 | 0.012325 | 10.5 | 0.63007 | 1.2059 | [5 empty] | -4.0188
10.6 | 0.012473 | -12.5579 | -4.0739 | 1.1724 |  |  |  | 0.53808 | 0.62825 | 12.4266 | 13.0549 | 2.5692 | 13.1075 | 0.012473 | 10.6 | 0.62825 | 1.1724 | [5 empty] | -4.1718
10.7 | 0.012956 | -12.9653 | -4.0402 | 1.1389 |  |  |  | 0.536 | 0.62643 | 12.9081 | 13.5345 | 2.6052 | 13.554 | 0.012956 | 10.7 | 0.62643 | 1.1389 | [5 empty] | -4.3157
10.8 | 0.013477 | -13.3693 | -4.5114 | 1.1054 |  |  |  | 0.53392 | 0.6246 | 13.4266 | 14.0512 | 2.6427 | 14.0157 | 0.013477 | 10.8 | 0.6246 | 1.1054 | [5 empty] | -4.4534
10.9 | 0.013811 | -13.8205 | -4.5452 | 1.0719 |  |  |  | 0.53184 | 0.62278 | 13.7598 | 14.3826 | 2.666 | 14.4932 | 0.013811 | 10.9 | 0.62278 | 1.0719 | [5 empty] | -4.6064
11 | 0.01422 | -14.275 | -4.8146 | 1.0384 |  |  |  | 0.52976 | 0.62095 | 14.1673 | 14.7882 | 2.6938 | 14.9869 | 0.01422 | 11 | 0.62095 | 1.0384 | [5 empty] | -4.7503
11.1 | 0.014629 | -14.7564 | -5.1175 | 1.0049 |  |  |  | 0.52768 | 0.61912 | 14.5746 | 15.1938 | 2.7209 | 15.4974 | 0.014629 | 11.1 | 0.61912 | 1.0049 | [5 empty] | -4.9156
11.2 | 0.015038 | -15.2682 | -5.2185 | 0.9714 |  |  |  | 0.5256 | 0.61729 | 14.982 | 15.5993 | 2.7472 | 16.0253 | 0.015038 | 11.2 | 0.61729 | 0.9714 | [5 empty] | -5.0931
11.3 | 0.015596 | -15.79 | -5.1849 | 0.9379 |  |  |  | 0.52352 | 0.61546 | 15.5375 | 16.153 | 2.7821 | 16.5711 | 0.015596 | 11.3 | 0.61546 | 0.9379 | [5 empty] | -5.2798
11.4 | 0.016376 | -16.3085 | -5.5217 | 0.9044 |  |  |  | 0.52144 | 0.61362 | 16.3152 | 16.9288 | 2.829 | 17.1356 | 0.016376 | 11.4 | 0.61362 | 0.9044 | [5 empty] | -5.4512
11.5 | 0.016711 | -16.8607 | -5.3869 | 0.87091 |  |  |  | 0.51936 | 0.61179 | 16.6486 | 17.2604 | 2.8484 | 17.7193 | 0.016711 | 11.5 | 0.61179 | 0.87091 | [5 empty] | -5.6318
11.6 | 0.01764 | -17.3994 | -5.6563 | 0.83741 |  |  |  | 0.51728 | 0.60995 | 17.5745 | 18.1844 | 2.9006 | 18.3229 | 0.01764 | 11.6 | 0.60995 | 0.83741 | [5 empty] | -5.7573
11.7 | 0.018124 | -17.965 | -6.0266 | 0.80391 |  |  |  | 0.5152 | 0.60811 | 18.0559 | 18.664 | 2.9266 | 18.9471 | 0.018124 | 11.7 | 0.60811 | 0.80391 | [5 empty] | -5.9226
11.8 | 0.018718 | -18.5677 | -6.094 | 0.77042 |  |  |  | 0.51312 | 0.60627 | 18.6485 | 19.2547 | 2.9578 | 19.5925 | 0.018718 | 11.8 | 0.60627 | 0.77042 | [5 empty] | -6.0511
11.9 | 0.019202 | -19.1771 | -6.3971 | 0.73692 |  |  |  | 0.51104 | 0.60442 | 19.13 | 19.7344 | 2.9824 | 20.2599 | 0.019202 | 11.9 | 0.60442 | 0.73692 | [5 empty] | -6.2164
12 | 0.019908 | -19.8168 | -6.5317 | 0.70342 |  |  |  | 0.50896 | 0.60258 | 19.8336 | 20.4362 | 3.0173 | 20.95 | 0.019908 | 12 | 0.60258 | 0.70342 | [5 empty] | -6.3756
12.1 | 0.020168 | -20.47 | -6.195 | 0.66993 |  |  |  | 0.50688 | 0.60073 | 20.0929 | 20.6936 | 3.0298 | 21.6636 | 0.020168 | 12.1 | 0.60073 | 0.66993 | [5 empty] | -6.596
12.2 | 0.020875 | -21.0895 | -6.9357 | 0.63643 |  |  |  | 0.5048 | 0.59888 | 20.7966 | 21.3955 | 3.0632 | 22.4016 | 0.020875 | 12.2 | 0.59888 | 0.63643 | [5 empty] | -6.8561
12.3 | 0.021692 | -21.783 | -6.6327 | 0.60294 |  |  |  | 0.50272 | 0.59703 | 21.6113 | 22.2083 | 3.1005 | 23.1646 | 0.021692 | 12.3 | 0.59703 | 0.60294 | [5 empty] | -7.104
12.4 | 0.022324 | -22.4463 | -7.003 | 0.56944 |  |  |  | 0.50064 | 0.59517 | 22.2409 | 22.8361 | 3.1283 | 23.9537 | 0.022324 | 12.4 | 0.59517 | 0.56944 | [5 empty] | -7.3673
12.5 | 0.02344 | -23.1466 | -7.2724 | 0.53594 |  |  |  | 0.49856 | 0.59332 | 23.352 | 23.9453 | 3.1758 | 24.7697 | 0.02344 | 12.5 | 0.59332 | 0.53594 | [5 empty] | -7.6152
12.6 | 0.023923 | -23.8738 | -7.8111 | 0.50245 |  |  |  | 0.49648 | 0.59146 | 23.8334 | 24.4249 | 3.1956 | 25.6134 | 0.023923 | 12.6 | 0.59146 | 0.50245 | [5 empty] | -7.9243
12.7 | 0.02448 | -24.6549 | -8.518 | 0.46895 |  |  |  | 0.4944 | 0.5896 | 24.3889 | 24.9785 | 3.218 | 26.4859 | 0.02448 | 12.7 | 0.5896 | 0.46895 | [5 empty] | -8.3069
12.8 | 0.025781 | -25.5067 | -8.7538 | 0.43545 |  |  |  | 0.49232 | 0.58774 | 25.6852 | 26.2729 | 3.2685 | 27.3881 | 0.025781 | 12.8 | 0.58774 | 0.43545 | [5 empty] | -8.5701
12.9 | 0.026042 | -26.3821 | -8.9895 | 0.40196 |  |  |  | 0.49024 | 0.58588 | 25.9444 | 26.5303 | 3.2783 | 28.3211 | 0.026042 | 12.9 | 0.58588 | 0.40196 | [5 empty] | -8.9711
13 | 0.026934 | -27.2811 | -9.1241 | 0.36846 |  |  |  | 0.48816 | 0.58401 | 26.8333 | 27.4173 | 3.3112 | 29.2858 | 0.026934 | 13 | 0.58401 | 0.36846 | [5 empty] | -9.3047
13.1 | 0.027938 | -28.1935 | -9.9321 | 0.33496 |  |  |  | 0.48608 | 0.58214 | 27.8332 | 28.4154 | 3.3469 | 30.2834 | 0.027938 | 13.1 | 0.58214 | 0.33496 | [5 empty] | -9.6444
13.2 | 0.028793 | -29.1867 | -10.4035 | 0.30147 |  |  |  | 0.484 | 0.58027 | 28.685 | 29.2653 | 3.3764 | 31.3149 | 0.028793 | 13.2 | 0.58027 | 0.30147 | [5 empty] | -9.9964
13.3 | 0.03028 | -30.227 | -9.8311 | 0.26797 |  |  |  | 0.48192 | 0.5784 | 30.1664 | 30.7448 | 3.4257 | 32.3816 | 0.03028 | 13.3 | 0.5784 | 0.26797 | [5 empty] | -10.2382
13.4 | 0.031358 | -31.2101 | -11.0431 | 0.23447 |  |  |  | 0.47984 | 0.57653 | 31.2404 | 31.817 |  | 33.4847 | 0.031358 | 13.4 | 0.57653 | 0.23447 | [5 empty] | -10.5259
13.5 | 0.03225 | -32.3145 | -10.6729 | 0.20098 |  |  |  | 0.47776 | 0.57465 | 32.1293 | 32.704 |  | 34.6253 | 0.03225 | 13.5 | 0.57465 | 0.20098 | [5 empty] | -10.7555
13.6 | 0.033514 | -33.3817 | -11.0095 | 0.16748 |  |  |  | 0.47568 | 0.57278 | 33.3885 | 33.9613 |  | 35.8047 | 0.033514 | 13.6 | 0.57278 | 0.16748 | [5 empty] | -10.9575
13.7 | 0.034889 | -34.4827 | -11.6828 | 0.13399 |  |  |  | 0.4736 | 0.5709 | 34.7588 | 35.3297 |  | 37.0244 | 0.034889 | 13.7 | 0.5709 | 0.13399 | [5 empty] | -11.1258
13.8 | 0.035967 | -35.651 | -11.1779 | 0.10049 |  |  |  | 0.47152 | 0.56902 | 35.8328 | 36.4018 |  | 38.2856 | 0.035967 | 13.8 | 0.56902 | 0.10049 | [5 empty] | -11.1809
13.9 | 0.036636 | -36.7688 | -11.9186 | 0.066993 |  |  |  | 0.46944 | 0.56713 | 36.4994 | 37.0666 |  | 39.5897 | 0.036636 | 13.9 | 0.56713 | 0.066993 | [5 empty] | -11.3493
14 | 0.038235 | -37.9606 | -11.5145 | 0.033496 |  |  |  | 0.46736 | 0.56525 | 38.0919 | 38.6571 |  | 40.9383 | 0.038235 | 14 | 0.56525 | 0.033496 | [5 empty] | -11.3585
14.1 | 0.038718 | -39.1121 | -11.3462 | 0 |  |  |  | 0.46528 | 0.56336 | 38.5734 | 39.1368 |  | 42.3328 | 0.038718 | 14.1 | 0.56336 | 0 | [5 empty] | -11.386
14.2 | 0.040093 | -40.2467 | -11.7839 | -0.033496 |  |  |  | 0.4632 | 0.56147 | 39.9437 | 40.5052 |  | 43.7748 | 0.040093 | 14.2 | 0.56147 | -0.033496 | [5 empty] | -11.3921
14.3 | 0.041692 | -41.4251 | -11.0095 | -0.066993 |  |  |  | 0.46112 | 0.55958 | 41.5361 | 42.0957 |  | 45.266 | 0.041692 | 14.3 | 0.55958 | -0.066993 | [5 empty] | -11.2575
14.4 | 0.042621 | -42.526 | -11.6829 | -0.10049 |  |  |  | 0.45904 | 0.55768 | 42.4621 | 43.0198 |  | 46.8079 | 0.042621 | 14.4 | 0.55768 | -0.10049 | [5 empty] | -11.1718
14.5 | 0.044517 | -43.6943 | -11.1443 | -0.13399 |  |  |  | 0.45696 | 0.55579 | 44.3509 | 44.9067 |  | 48.4023 | 0.044517 | 14.5 | 0.55579 | -0.13399 | [5 empty] | -10.9361
14.6 | 0.044963 | -44.8088 | -10.9758 | -0.16748 |  |  |  | 0.45488 | 0.55389 | 44.7953 | 45.3492 |  | 50.0511 | 0.044963 | 14.6 | 0.55389 | -0.16748 | [5 empty] | -10.7188
14.7 | 0.046041 | -45.9063 | -11.0769 | -0.20098 |  |  |  | 0.4528 | 0.55199 | 45.8693 | 46.4213 |  | 51.756 | 0.046041 | 14.7 | 0.55199 | -0.20098 | [5 empty] | -10.4892
14.8 | 0.0479 | -47.014 | -10.2015 | -0.23447 |  |  |  | 0.45072 | 0.55009 | 47.721 | 48.2711 |  | 53.519 | 0.0479 | 14.8 | 0.55009 | -0.23447 | [5 empty] | -10.1219
14.9 | 0.048123 | -48.0342 | -10.2352 | -0.26797 |  |  |  | 0.44864 | 0.54818 | 47.9433 | 48.4915 |  | 55.3421 | 0.048123 | 14.9 | 0.54818 | -0.26797 | [5 empty] | -9.8556
15 | 0.049536 | -49.0577 | -9.3261 | -0.30147 |  |  |  | 0.44656 | 0.54627 | 49.3506 | 49.8969 |  | 57.2273 | 0.049536 | 15 | 0.54627 | -0.30147 | [5 empty] | -9.4669
15.1 | 0.050539 | -49.9903 | -9.1241 | -0.33496 |  |  |  | 0.44448 | 0.54436 | 50.3506 | 50.895 |  | 59.1766 | 0.050539 | 15.1 | 0.54436 | -0.33496 | [5 empty] | -9.121
15.2 | 0.050837 | -50.9027 | -8.8211 | -0.36846 |  |  |  | 0.4424 | 0.54245 | 50.6468 | 51.1892 |  | 61.1924 | 0.050837 | 15.2 | 0.54245 | -0.36846 | [5 empty] | -8.8089
15.3 | 0.052324 | -51.7848 | -7.7438 | -0.40196 |  |  |  | 0.44032 | 0.54054 | 52.1283 | 52.6689 |  | 63.2769 | 0.052324 | 15.3 | 0.54054 | -0.40196 | [5 empty] | -8.4018
15.4 | 0.052956 | -52.5592 | -8.0803 | -0.43545 |  |  |  | 0.43824 | 0.53862 | 52.7578 | 53.2964 |  | 65.4323 | 0.052956 | 15.4 | 0.53862 | -0.43545 | [5 empty] | -8.059
15.5 | 0.053922 | -53.3672 | -7.407 | -0.46895 |  |  |  | 0.43616 | 0.5367 | 53.7207 | 54.2575 |  | 67.6612 | 0.053922 | 15.5 | 0.5367 | -0.46895 | [5 empty] | -7.6856
15.6 | 0.054814 | -54.1079 | -7.3397 | -0.50245 |  |  |  | 0.43408 | 0.53478 | 54.6096 | 55.1444 |  | 69.966 | 0.054814 | 15.6 | 0.53478 | -0.50245 | [5 empty] | -7.3856
15.7 | 0.055037 | -54.8419 | -7.5417 | -0.53594 |  |  |  | 0.432 | 0.53286 | 54.8318 | 55.3646 |  | 72.3493 | 0.055037 | 15.7 | 0.53286 | -0.53594 | [5 empty] | -7.1438
15.8 | 0.055781 | -55.5961 | -6.5989 | -0.56944 |  |  |  | 0.42992 | 0.53094 | 55.5725 | 56.1034 |  | 74.8138 | 0.055781 | 15.8 | 0.53094 | -0.56944 | [5 empty] | -6.8775
15.9 | 0.05645 | -56.256 | -6.4307 | -0.60294 |  |  |  | 0.42784 | 0.52901 | 56.2392 | 56.7682 |  | 77.3622 | 0.05645 | 15.9 | 0.52901 | -0.60294 | [5 empty] | -6.7092
16 | 0.057045 | -56.899 | -6.1277 | -0.63643 |  |  |  | 0.42576 | 0.52708 | 56.8317 | 57.3588 |  | 79.9974 | 0.057045 | 16 | 0.52708 | -0.63643 | [5 empty] | -6.5041
16.1 | 0.057714 | -57.5118 | -6.0267 | -0.66993 |  |  |  | 0.42368 | 0.52515 | 57.4984 | 58.0235 |  | 82.7224 | 0.057714 | 16.1 | 0.52515 | -0.66993 | [5 empty] | -6.3205
16.2 | 0.058643 | -58.1145 | -6.4643 | -0.70342 |  |  |  | 0.4216 | 0.52321 | 58.4242 | 58.9474 |  | 85.5403 | 0.058643 | 16.2 | 0.52321 | -0.70342 | [5 empty] | -6.1368
16.3 | 0.059164 | -58.7609 | -5.892 | -0.73692 |  |  |  | 0.41952 | 0.52127 | 58.9427 | 59.4639 |  | 88.4541 | 0.059164 | 16.3 | 0.52127 | -0.73692 | [5 empty] | -5.9287
16.4 | 0.05961 | -59.3501 | -5.8919 | -0.77042 |  |  |  | 0.41744 | 0.51933 | 59.3871 | 59.9064 |  | 91.4672 | 0.05961 | 16.4 | 0.51933 | -0.77042 | [5 empty] | -5.8093
16.5 | 0.060056 | -59.9393 | -5.8247 | -0.80391 |  |  |  | 0.41536 | 0.51739 | 59.8315 | 60.3489 |  | 94.5829 | 0.060056 | 16.5 | 0.51739 | -0.80391 | [5 empty] | -5.7114
16.6 | 0.060688 | -60.5218 | -5.3869 | -0.83741 |  |  |  | 0.41328 | 0.51545 | 60.4612 | 60.9766 |  | 97.8047 | 0.060688 | 16.6 | 0.51545 | -0.83741 | [5 empty] | -5.6226
16.7 | 0.061469 | -61.0605 | -5.3196 | -0.87091 |  |  |  | 0.4112 | 0.5135 | 61.239 | 61.7525 |  | 101.1363 | 0.061469 | 16.7 | 0.5135 | -0.87091 | [5 empty] | -5.5002
16.8 | 0.062175 | -61.5924 | -5.2523 | -0.9044 |  |  |  | 0.40912 | 0.51155 | 61.9425 | 62.4541 |  | 104.5814 | 0.062175 | 16.8 | 0.51155 | -0.9044 | [5 empty] | -5.3288
16.9 | 0.062286 | -62.1176 | -5.2859 | -0.9379 |  |  |  | 0.40704 | 0.5096 | 62.0537 | 62.5633 |  | 108.1438 | 0.062286 | 16.9 | 0.5096 | -0.9379 | [5 empty] | -5.2125
17 | 0.062955 | -62.6462 | -5.3533 | -0.9714 |  |  |  | 0.40496 | 0.50765 | 62.7203 | 63.2279 |  | 111.8276 | 0.062955 | 17 | 0.50765 | -0.9714 | [5 empty] | -5.0809
17.1 | 0.063476 | -63.1816 | -5.1512 | -1.0049 |  |  |  | 0.40288 | 0.50569 | 63.2389 | 63.7445 |  | 115.6369 | 0.063476 | 17.1 | 0.50569 | -1.0049 | [5 empty] | -4.9462
17.2 | 0.063662 | -63.6967 | -4.6798 | -1.0384 |  |  |  | 0.4008 | 0.50373 | 63.424 | 63.9277 |  | 119.5759 | 0.063662 | 17.2 | 0.50373 | -1.0384 | [5 empty] | -4.8758
17.3 | 0.064517 | -64.1647 | -4.5789 | -1.0719 |  |  |  | 0.39872 | 0.50177 | 64.2758 | 64.7775 |  | 123.6491 | 0.064517 | 17.3 | 0.50177 | -1.0719 | [5 empty] | -4.8023
17.4 | 0.064963 | -64.6226 | -4.6125 | -1.1054 |  |  |  | 0.39664 | 0.4998 | 64.7202 | 65.22 |  | 127.8611 | 0.064963 | 17.4 | 0.4998 | -1.1054 | [5 empty] | -4.7135
17.5 | 0.065446 | -65.0838 | -4.4442 | -1.1389 |  |  |  | 0.39456 | 0.49784 | 65.2016 | 65.6994 |  | 132.2165 | 0.065446 | 17.5 | 0.49784 | -1.1389 | [5 empty] | -4.6126
17.6 | 0.065967 | -65.5282 | -4.3432 | -1.1724 |  |  |  | 0.39248 | 0.49587 | 65.7201 | 66.216 |  | 136.7203 | 0.065967 | 17.6 | 0.49587 | -1.1724 | [5 empty] | -4.533
17.7 | 0.066376 | -65.9625 | -4.6125 | -1.2059 |  |  |  | 0.3904 | 0.4939 | 66.1275 | 66.6214 |  | 141.3775 | 0.066376 | 17.7 | 0.4939 | -1.2059 | [5 empty] | -4.4411
17.8 | 0.066636 | -66.4238 | -4.5116 | -1.2394 |  |  |  | 0.38832 | 0.49192 | 66.3867 | 66.8787 |  | 146.1933 | 0.066636 | 17.8 | 0.49192 | -1.2394 | [5 empty] | -4.4197
17.9 | 0.06723 | -66.875 | -4.2758 | -1.2729 |  |  |  | 0.38624 | 0.48994 | 66.9793 | 67.4693 |  | 151.1732 | 0.06723 | 17.9 | 0.48994 | -1.2729 | [5 empty] | -4.4044
18 | 0.067379 | -67.3025 | -4.175 | -1.3064 |  |  |  | 0.38416 | 0.48796 | 67.1275 | 67.6154 |  | 156.3228 | 0.067379 | 18 | 0.48796 | -1.3064 | [5 empty] | -4.3922
18.1 | 0.067863 | -67.72 | -4.4778 | -1.3399 |  |  |  | 0.38208 | 0.48598 | 67.609 | 68.0949 |  | 161.6477 | 0.067863 | 18.1 | 0.48598 | -1.3399 | [5 empty] | -4.3891
18.2 | 0.068271 | -68.1678 | -4.1412 | -1.3734 |  |  |  | 0.38 | 0.484 | 68.0163 | 68.5003 |  | 167.154 | 0.068271 | 18.2 | 0.484 | -1.3734 | [5 empty] | -4.3952
18.3 | 0.068755 | -68.5819 | -4.4442 | -1.4068 |  |  |  | 0.37792 | 0.48201 | 68.4977 | 68.9797 |  | 172.8479 | 0.068755 | 18.3 | 0.48201 | -1.4068 | [5 empty] | -4.4014
18.4 | 0.069498 | -69.0264 | -4.4106 | -1.4403 |  |  |  | 0.37584 | 0.48002 | 69.2386 | 69.7186 |  | 178.7358 | 0.069498 | 18.4 | 0.48002 | -1.4403 | [5 empty] | -4.383
18.5 | 0.069684 | -69.4674 | -4.4778 | -1.4738 |  |  |  | 0.37376 | 0.47802 | 69.4236 | 69.9016 |  | 184.8242 | 0.069684 | 18.5 | 0.47802 | -1.4738 | [5 empty] | -4.4136
18.6 | 0.070056 | -69.9152 | -4.4106 | -1.5073 |  |  |  | 0.37168 | 0.47603 | 69.7941 | 70.2701 |  | 191.12 | 0.070056 | 18.6 | 0.47603 | -1.5073 | [5 empty] | -4.4656
18.7 | 0.070911 | -70.3563 | -4.4105 | -1.5408 |  |  |  | 0.36961 | 0.47403 | 70.6458 | 71.1198 |  | 197.6302 | 0.070911 | 18.7 | 0.47403 | -1.5408 | [5 empty] | -4.4258
18.8 | 0.070948 | -70.7973 | -4.6799 | -1.5743 |  |  |  | 0.36753 | 0.47203 | 70.6828 | 71.1549 |  | 204.3623 | 0.070948 | 18.8 | 0.47203 | -1.5743 | [5 empty] | -4.4595
18.9 | 0.071543 | -71.2653 | -4.3095 | -1.6078 |  |  |  | 0.36545 | 0.47002 | 71.2754 | 71.7454 |  | 211.3236 | 0.071543 | 18.9 | 0.47002 | -1.6078 | [5 empty] | -4.4442
19 | 0.0721 | -71.6962 | -4.6126 | -1.6413 |  |  |  | 0.36337 | 0.46801 | 71.8309 | 72.2989 |  | 218.5221 | 0.0721 | 19 | 0.46801 | -1.6413 | [5 empty] | -4.4197
19.1 | 0.072323 | -72.1575 | -4.7472 | -1.6748 |  |  |  | 0.36129 | 0.466 | 72.0531 | 72.5191 |  | 225.9657 | 0.072323 | 19.1 | 0.466 | -1.6748 | [5 empty] | -4.4381
19.2 | 0.072732 | -72.6322 | -4.0402 | -1.7083 |  |  |  | 0.35921 | 0.46399 | 72.4606 | 72.9246 |  | 233.663 | 0.072732 | 19.2 | 0.46399 | -1.7083 | [5 empty] | -4.4564
19.3 | 0.073141 | -73.0363 | -4.5115 | -1.7418 |  |  |  | 0.35713 | 0.46197 | 72.8678 | 73.3298 |  | 241.6224 | 0.073141 | 19.3 | 0.46197 | -1.7418 | [5 empty] | -4.4656
19.4 | 0.073922 | -73.4874 | -4.2758 | -1.7753 |  |  |  | 0.35505 | 0.45995 | 73.6456 | 74.1056 |  | 249.853 | 0.073922 | 19.4 | 0.45995 | -1.7753 | [5 empty] | -4.4442
19.5 | 0.074256 | -73.915 | -4.1412 | -1.8088 |  |  |  | 0.35297 | 0.45793 | 73.979 | 74.4369 |  | 258.3639 | 0.074256 | 19.5 | 0.45793 | -1.8088 | [5 empty] | -4.4412
19.6 | 0.074777 | -74.3291 | -4.6799 | -1.8423 |  |  |  | 0.35089 | 0.45591 | 74.4974 | 74.9533 |  | 267.1647 | 0.074777 | 19.6 | 0.45591 | -1.8423 | [5 empty] | -4.4044
19.7 | 0.075297 | -74.7971 | -4.6125 | -1.8758 |  |  |  | 0.34881 | 0.45388 | 75.016 | 75.4699 |  | 276.2654 | 0.075297 | 19.7 | 0.45388 | -1.8758 | [5 empty] | -4.3738
19.8 | 0.075372 | -75.2583 | -4.5116 | -1.9093 |  |  |  | 0.34673 | 0.45185 | 75.09 | 75.5419 |  | 285.676 | 0.075372 | 19.8 | 0.45185 | -1.9093 | [5 empty] | -4.4075
19.9 | 0.075929 | -75.7095 | -4.4442 | -1.9428 |  |  |  | 0.34465 | 0.44981 | 75.6455 | 76.0953 |  | 295.4072 | 0.075929 | 19.9 | 0.44981 | -1.9428 | [5 empty] | -4.383
20 | 0.076264 | -76.1539 | -4.2759 | -1.9763 |  |  |  | 0.34257 | 0.44777 | 75.9788 | 76.4266 |  | 305.4699 | 0.076264 | 20 | 0.44777 | -1.9763 | [5 empty] | -4.4136
20.1 | 0.076673 | -76.5815 | -4.2086 | -2.0098 |  |  |  | 0.34049 | 0.44573 | 76.3863 | 76.832 |  | 315.8753 | 0.076673 | 20.1 | 0.44573 | -2.0098 | [5 empty] | -4.4534
20.2 | 0.077491 | -77.0024 | -4.4105 | -2.0433 |  |  |  | 0.33841 | 0.44369 | 77.201 | 77.6447 |  | 326.6352 | 0.077491 | 20.2 | 0.44369 | -2.0433 | [5 empty] | -4.4136
20.3 | 0.077825 | -77.4434 | -4.4106 | -2.0768 |  |  |  | 0.33633 | 0.44164 | 77.5343 | 77.976 |  | 337.7616 | 0.077825 | 20.3 | 0.44164 | -2.0768 | [5 empty] | -4.3983
20.4 | 0.078122 | -77.8845 | -4.2422 | -2.1103 |  |  |  | 0.33425 | 0.43959 | 77.8306 | 78.2702 |  | 349.2671 | 0.078122 | 20.4 | 0.43959 | -2.1103 | [5 empty] | -4.3708
20.5 | 0.078829 | -78.3087 | -4.6126 | -2.1438 |  |  |  | 0.33217 | 0.43754 | 78.5343 | 78.9718 |  | 361.1644 | 0.078829 | 20.5 | 0.43754 | -2.1438 | [5 empty] | -4.3279
20.6 | 0.078978 | -78.77 | -4.5788 | -2.1773 |  |  |  | 0.33009 | 0.43548 | 78.6824 | 79.1179 |  | 373.467 | 0.078978 | 20.6 | 0.43548 | -2.1773 | [5 empty] | -4.3187
20.7 | 0.079424 | -79.2278 | -4.2422 | -2.2108 |  |  |  | 0.32801 | 0.43342 | 79.1269 | 79.5603 |  | 386.1887 | 0.079424 | 20.7 | 0.43342 | -2.2108 | [5 empty] | -4.3126
20.8 | 0.080167 | -79.6521 | -4.4442 | -2.2443 |  |  |  | 0.32593 | 0.43136 | 79.8675 | 80.2989 |  | 399.3437 | 0.080167 | 20.8 | 0.43136 | -2.2443 | [5 empty] | -4.2881
20.9 | 0.080241 | -80.0965 | -4.2086 | -2.2778 |  |  |  | 0.32385 | 0.42929 | 79.9416 | 80.3709 |  | 412.9469 | 0.080241 | 20.9 | 0.42929 | -2.2778 | [5 empty] | -4.2851
21 | 0.080613 | -80.5173 | -3.9729 | -2.3113 |  |  |  | 0.32177 | 0.42722 | 80.3119 | 80.7392 |  | 427.0134 | 0.080613 | 21 | 0.42722 | -2.3113 | [5 empty] | -4.2789
21.1 | 0.081357 | -80.9146 | -4.1749 | -2.3447 |  |  |  | 0.31969 | 0.42515 | 81.0527 | 81.4778 |  | 441.559 | 0.081357 | 21.1 | 0.42515 | -2.3447 | [5 empty] | -4.2392
21.2 | 0.081728 | -81.3321 | -4.1412 | -2.3782 |  |  |  | 0.31761 | 0.42308 | 81.423 | 81.846 |  | 456.6002 | 0.081728 | 21.2 | 0.42308 | -2.3782 | [5 empty] | -4.1902
21.3 | 0.082174 | -81.7462 | -4.1412 | -2.4117 |  |  |  | 0.31553 | 0.421 | 81.8674 | 82.2884 |  | 472.1537 | 0.082174 | 21.3 | 0.421 | -2.4117 | [5 empty] | -4.1565
21.4 | 0.082732 | -82.1604 | -4.3769 | -2.4452 |  |  |  | 0.31345 | 0.41891 | 82.4229 | 82.8418 |  | 488.237 | 0.082732 | 21.4 | 0.41891 | -2.4452 | [5 empty] | -4.1137
21.5 | 0.082769 | -82.598 | -4.175 | -2.4787 |  |  |  | 0.31137 | 0.41683 | 82.4601 | 82.8769 |  | 504.8682 | 0.082769 | 21.5 | 0.41683 | -2.4787 | [5 empty] | -4.1014
21.6 | 0.083215 | -83.0155 | -4.1749 | -2.5122 |  |  |  | 0.30929 | 0.41474 | 82.9045 | 83.3192 |  | 522.0659 | 0.083215 | 21.6 | 0.41474 | -2.5122 | [5 empty] | -4.1045
21.7 | 0.083587 | -83.433 | -4.0402 | -2.5457 |  |  |  | 0.30721 | 0.41264 | 83.2748 | 83.6875 |  | 539.8494 | 0.083587 | 21.7 | 0.41264 | -2.5457 | [5 empty] | -4.1075
21.8 | 0.083996 | -83.837 | -3.8719 | -2.5792 |  |  |  | 0.30513 | 0.41055 | 83.6822 | 84.0927 |  | 558.2387 | 0.083996 | 21.8 | 0.41055 | -2.5792 | [5 empty] | -4.1045
21.9 | 0.084739 | -84.2242 | -3.9729 | -2.6127 |  |  |  | 0.30305 | 0.40845 | 84.4228 | 84.8313 |  | 577.2544 | 0.084739 | 21.9 | 0.40845 | -2.6127 | [5 empty] | -4.0953
22 | 0.085074 | -84.6215 | -4.0738 | -2.6462 |  |  |  | 0.30097 | 0.40634 | 84.7562 | 85.1625 |  | 596.9178 | 0.085074 | 22 | 0.40634 | -2.6462 | [5 empty] | -4.0616
22.1 | 0.085223 | -85.0289 | -4.0064 | -2.6797 |  |  |  | 0.29889 | 0.40423 | 84.9044 | 85.3086 |  | 617.2511 | 0.085223 | 22.1 | 0.40423 | -2.6797 | [5 empty] | -4.0433
22.2 | 0.085966 | -85.4296 | -4.2086 | -2.7132 |  |  |  | 0.29681 | 0.40212 | 85.645 | 86.0472 |  | 638.277 | 0.085966 | 22.2 | 0.40212 | -2.7132 | [5 empty] | -4.0157
22.3 | 0.086189 | -85.8504 | -4.1075 | -2.7467 |  |  |  | 0.29473 | 0.40001 | 85.8672 | 86.2672 |  | 660.0191 | 0.086189 | 22.3 | 0.40001 | -2.7467 | [5 empty] | -3.9912
22.4 | 0.086449 | -86.2612 | -4.0402 | -2.7802 |  |  |  | 0.29265 | 0.39789 | 86.1265 | 86.5244 |  | 682.5018 | 0.086449 | 22.4 | 0.39789 | -2.7802 | [5 empty] | -3.9943
22.5 | 0.087119 | -86.6652 | -4.0065 | -2.8137 |  |  |  | 0.29057 | 0.39576 | 86.7931 | 87.1889 |  | 705.7504 | 0.087119 | 22.5 | 0.39576 | -2.8137 | [5 empty] | -3.982
22.6 | 0.087267 | -87.0658 | -3.9728 | -2.8472 |  |  |  | 0.28849 | 0.39364 | 86.9413 | 87.3349 |  | 729.7909 | 0.087267 | 22.6 | 0.39364 | -2.8472 | [5 empty] | -3.9729
22.7 | 0.087639 | -87.4631 | -3.8718 | -2.8807 |  |  |  | 0.28641 | 0.39151 | 87.3116 | 87.7031 |  | 754.6503 | 0.087639 | 22.7 | 0.39151 | -2.8807 | [5 empty] | -3.9912
22.8 | 0.088234 | -87.8503 | -3.7709 | -2.9142 |  |  |  | 0.28433 | 0.38937 | 87.9043 | 88.2936 |  | 780.3565 | 0.088234 | 22.8 | 0.38937 | -2.9142 | [5 empty] | -3.9667
22.9 | 0.088531 | -88.2274 | -3.9054 | -2.9477 |  |  |  | 0.28225 | 0.38724 | 88.2004 | 88.5877 |  | 806.9383 | 0.088531 | 22.9 | 0.38724 | -2.9477 | [5 empty] | -3.9361
23 | 0.0892 | -88.6179 | -3.8382 | -2.9812 |  |  |  | 0.28017 | 0.38509 | 88.867 | 89.2521 |  | 834.4257 | 0.0892 | 23 | 0.38509 | -2.9812 | [5 empty] | -3.9055
23.1 | 0.089498 | -89.0018 | -3.9728 | -3.0147 |  |  |  | 0.27809 | 0.38295 | 89.1633 | 89.5463 |  | 862.8493 | 0.089498 | 23.1 | 0.38295 | -3.0147 | [5 empty] | -3.8657
23.2 | 0.089609 | -89.3991 | -4.2086 | -3.0482 |  |  |  | 0.27601 | 0.3808 | 89.2745 | 89.6553 |  | 892.2412 | 0.089609 | 23.2 | 0.3808 | -3.0482 | [5 empty] | -3.8657
23.3 | 0.090241 | -89.8199 | -3.9391 | -3.0817 |  |  |  | 0.27393 | 0.37864 | 89.9041 | 90.2827 |  | 922.6343 | 0.090241 | 23.3 | 0.37864 | -3.0817 | [5 empty] | -3.8565
23.4 | 0.090353 | -90.2138 | -3.7709 | -3.1152 |  |  |  | 0.27185 | 0.37648 | 90.0152 | 90.3917 |  | 954.0626 | 0.090353 | 23.4 | 0.37648 | -3.1152 | [5 empty] | -3.8596
23.5 | 0.090762 | -90.5909 | -3.7036 | -3.1487 |  |  |  | 0.26977 | 0.37432 | 90.4225 | 90.7968 |  | 986.5616 | 0.090762 | 23.5 | 0.37432 | -3.1487 | [5 empty] | -3.8688
23.6 | 0.091356 | -90.9613 | -3.5688 | -3.1822 |  |  |  | 0.26769 | 0.37216 | 91.0152 | 91.3873 |  | 1020.1675 | 0.091356 | 23.6 | 0.37216 | -3.1822 | [5 empty] | -3.8473
23.7 | 0.091654 | -91.3181 | -3.9727 | -3.2157 |  |  |  | 0.26561 | 0.36998 | 91.3114 | 91.6814 |  | 1054.9183 | 0.091654 | 23.7 | 0.36998 | -3.2157 | [5 empty] | -3.8167
23.8 | 0.092286 | -91.7154 | -3.7709 | -3.2492 |  |  |  | 0.26353 | 0.36781 | 91.941 | 92.3088 |  | 1090.8527 | 0.092286 | 23.8 | 0.36781 | -3.2492 | [5 empty] | -3.7678
23.9 | 0.092583 | -92.0925 | -3.8045 | -3.2826 |  |  |  | 0.26145 | 0.36563 | 92.2373 | 92.6029 |  | 1128.0112 | 0.092583 | 23.9 | 0.36563 | -3.2826
24 | 0.092695 | -92.473 | -4.0066 | -3.3161 |  |  |  | 0.25937 | 0.36345 | 92.3484 | 92.7118 |  | 1166.4355 | 0.092695 | 24 | 0.36345 | -3.3161
24.1 | 0.093289 | -92.8736 | -3.6024 | -3.3496 |  |  |  | 0.25729 | 0.36126 | 92.941 | 93.3022 |  | 1206.1687 | 0.093289 | 24.1 | 0.36126 | -3.3496
24.2 | 0.093438 | -93.2339 | -3.6362 | -3.3831 |  |  |  | 0.25521 | 0.35907 | 93.089 | 93.4481 |  | 1247.2553 | 0.093438 | 24.2 | 0.35907 | -3.3831
24.3 | 0.093996 | -93.5975 | -3.6698 | -3.4166 |  |  |  | 0.25313 | 0.35687 | 93.6445 | 94.0014 |  | 1289.7415 | 0.093996 | 24.3 | 0.35687 | -3.4166
24.4 | 0.094405 | -93.9645 |  | -3.4501 |  |  |  | 0.25105 | 0.35467 | 94.052 | 94.4067 |  | 1333.6749 | 0.094405 | 24.4 | 0.35467 | -3.4501
Floating potential method Rp: 24.5; Floating potential method L: 0.094553; Floating potential method Gas: -3.4836; Type into YELLOW cells only xp: 0.24897; Type into YELLOW cells only p: 0.35246 | 94.2001 | 94.5526 |  | 1379.1048 | 0.094553 | 24.5 | 0.35246 | -3.4836
Floating potential method Rp: 24.6; Floating potential method L: 0.095185; Floating potential method Gas: -3.5171; Type into YELLOW cells only xp: 0.24689; Type into YELLOW cells only p: 0.35025 | 94.8298 | 95.18 |  | 1426.0823 | 0.095185 | 24.6 | 0.35025 | -3.5171
Floating potential method Rp: 24.7; Floating potential method L: 0.095334; Floating potential method Gas: -3.5506; Type into YELLOW cells only xp: 0.24481; Type into YELLOW cells only p: 0.34804 | 94.9778 | 95.3259 |  | 1474.66 | 0.095334 | 24.7 | 0.34804 | -3.5506
Floating potential method Rp: 24.8; Floating potential method L: 0.095669; Floating potential method Gas: -3.5841; Type into YELLOW cells only xp: 0.24273; Type into YELLOW cells only p: 0.34582 | 95.3112 | 95.657 |  | 1524.8924 | 0.095669 | 24.8 | 0.34582 | -3.5841
Floating potential method Rp: 24.9; Floating potential method L: 0.096338; Floating potential method Gas: -3.6176; Type into YELLOW cells only xp: 0.24065; Type into YELLOW cells only p: 0.34359 | 95.9778 | 96.3214 |  | 1576.836 | 0.096338 | 24.9 | 0.34359 | -3.6176
Note on H3: This is for the CL case.
Note on H8: This uses measured Vf, not theor. one.
Note on I3: This is for the CL case.
Note on R9: The value of xp used here is for the CL density, not the ABR density.

## Program Notes
About the program...
First, fill in the yellow boxes:
1 | The probe radius Rp and length L, and the atomic number of the gas.
2 | The p and a0 boxes are normally not changed from 4/3 and exp(-1/2).
However, one can experiment with other values if desired.
3 | "Paste Value" the probe data into columns A and B, with Vp in volts and I in amps, ion current being positive.
All the other values are calculated automatically, and the results are in the Output box.
The space potential Vs is determined by the voltage at which dI/dV is minimum. Even after digital noise has been removed,
the curve is still subject to plasma noise. dI/dV is further smoothed to remove this and force the creation of a
minimum. This minimum by found by the MIN function, and its location by the MATCH function. Vp at this
location is then called Vs, and is found with the INDIRECT function.
The ion current is raised to the power p, and the resulting curve is fitted with a straight line with the use of
the SLOPE and INTERCEPT functions of Excel.
The fit is made between Vmin and Vmax. Vmin is the first data point + 2 volts, to eliminate jitter at the beginning
of the trace; and Vmax is Vs - 4KTe, where the contribution of a Maxwellian electron distribution is <2%.
Since Te will be affected by this fit, it cannot be used to evaluate Vmax. A rough estimate is used instead.
The difference Vs - Vf is about 4KTe for a cylindrical probe. Thus, KTe can be estimated from Vf, whose
measurement is described later.
The extrapolated ion current is then subtracted from the electron current, and the result is plotted on a semilog scale.
A straight line is fitted to this curve with the SLOPE and INTERCEPT functions. The fit is made between
Vmin and Vmax. Vmax is just short of Vs, say, Vs-KTe/4, where there should be almost no curvature.
Vmin is 4 e-foldings down, at Vmax-4KTe. Again the rough estimate of KTe is used.
An accurate value of Te is given by the reciprocal of this straight line. The magniture of Ie at Vs gives the electron
density. Of course, it is not accurate, since it depends exponentially on the estimate of Vs.
The floating potential Vf, at which I = 0, is found with the VLOOKUP function. At the same time, the extrapolated ion
current If at Vf and the value of hf =(Vs-Vf)/KTe is found the same way.
The values of If, Vf, and Te are used to calculate the ion density Ni according to the Child-Langmuir formula.
Since this formula involves the Debye length, which depends on root Ni, a quadratic equation is solved to get Ni(CL).
The measured If and Te also can be used to calculate Ni using the exact ABR formulation. The ABR curves
parametrized by Chen and Arnush are used. However, this formula is nonlinear and must be solved by iteration.
An automatic iteration procedure is used to give Ni(ABR).
The entire calculation for a Hiden ESP data set of 1000 points requires only about 0.1 sec on a modern PC.